Regulation (EU) No 528/2012 concerning the making available on the market and use of biocidal products

PRODUCT ASSESSMENT REPORT OF A BIOCIDAL PRODUCT (FAMILY) FOR NATIONAL AUTHORISATION APPLICATIONS

(submitted by the evaluating Competent Authority)

HYDROGEN PEROXIDE SOLUTION

Product type(s) 2

HYDROGEN PEROXIDE as included in the Union list of approved active substances

Case Number in R4BP: BC-RJ029431-39

Evaluating Competent Authority: France

Date: July 2019

Table of Contents

Table of Contents	2

1	CONCLUSION	4

2	ASSESSMENT REPORT	6

2.1	Summary of the product assessment	6

2.1.1	Administrative information	6

2.1.1.1	Identifier of the product / product family	6

2.1.1.2	Authorisation holder	6

2.1.1.3	Manufacturer(s) of the products of the family	6

2.1.1.4	Manufacturer(s) of the active substance(s)	6

2.1.2	Product (family) composition and formulation	6

2.1.2.1	Identity of the active substance	7

2.1.2.2	Candidate(s) for substitution	7

2.1.2.3	Qualitative and quantitative information on the composition of the biocidal product	7

2.1.2.4	Qualitative and quantitative information on the composition of the biocidal product family	7

2.1.2.5	Information on technical equivalence	7

2.1.2.6	Information on the substance(s) of concern	7

2.1.2.7	Type of formulation	7

2.1.3	Hazard and precautionary statements	8

2.1.4	Authorised use(s)	9

2.1.4.1	Use description	9

2.1.4.2	Use-specific instructions for use	10

2.1.4.3	Use-specific risk mitigation measures	10

2.1.4.4	Where specific to the use, the particulars of likely direct or indirect effects, first aid instructions and emergency measures to protect the environment	10

2.1.4.5	Where specific to the use, the instructions for safe disposal of the product and its packaging	10

2.1.4.6	Where specific to the use, the conditions of storage and shelf-life of the product under normal conditions of storage	10

2.1.5	General directions for use	10

2.1.5.1	Instructions for use	10

2.1.5.2	Risk mitigation measures	11

2.1.5.3	Particulars of likely direct or indirect effects, first aid instructions and emergency measures to protect the environment	11

2.1.5.4	Instructions for safe disposal of the product and its packaging	12

2.1.5.5	Conditions of storage and shelf-life of the product under normal conditions of storage	12

2.1.6	Other information	12

2.1.7	Packaging of the biocidal product	12

2.1.8	Documentation	12

2.1.8.1	Data submitted in relation to product application	12

2.1.8.2	Access to documentation	13

2.2	Assessment of the biocidal product	13

2.2.1	Intended use(s) as applied for by the applicant	13

2.2.2	Identity	15

2.2.3	Physical, chemical and technical properties	16

2.2.4	Physical hazards and respective characteristics	23

2.2.5	Methods for detection and identification	26

2.2.6	Efficacy against target organisms	29

2.2.6.1	Function and field of use	29

2.2.6.2	Organisms to be controlled and products, organisms or objects to be protected	29

2.2.6.3	Effects on target organisms, including unacceptable suffering	30

2.2.6.4	Mode of action, including time delay	30

2.2.6.5	Efficacy data	30

2.2.6.6	Occurrence of resistance and resistance management	45

2.2.6.7	Known limitations	46

2.2.6.8	Evaluation of the label claims	46

2.2.6.9	Relevant information if the product is intended to be authorised for use with other biocidal product(s)	46

2.2.7	Risk assessment for human health	47

2.2.7.1	Assessment of effects on Human Health	47

2.2.7.2	Exposure assessment	50

2.2.7.3	Risk characterisation for human health	59

2.2.8	Risk assessment for animal health	63

2.2.9	Risk assessment for the environment	63

2.2.9.1	Effects assessment on the environment	63

2.2.9.2	Exposure assessment	65

2.2.9.3	Risk characterisation	66

2.2.10	Measures to protect man, animals and the environment	67

2.2.11	Assessment of a combination of biocidal products	67

2.2.12	Comparative assessment	67

3	Annexes	69

3.1	List of studies for the biocidal product	69

3.2	Output tables from exposure assessment tools	69

3.3	New information on the active substance	69

3.4	Residue behaviour	70

3.5	Summaries of the efficacy studies	70

3.6	Confidential annex	70

# CONCLUSION

HYDROGEN PEROXIDE SOLUTION containing hydrogen peroxide is a PT2 biocidal product, intended for the disinfection of instruments and materials in the industrial sectors, against bacteria, bacterial spores, yeasts, fungi and viruses. The product is applied by vaporising hydrogen peroxide in the STERRAD® 200 GMP System process by professional user only.

In the frame of application for marketing authorisation of HYDROGEN PEROXIDE SOLUTION as a biocidal product, the use for the disinfection of instruments and materials which is covered by the medical devices Regulation is not concerned.

Physico-chemical properties

Physico-chemical studies were provided for HYDROGEN PEROXIDE SOLUTION. A storage stability (12 months at 30°C) study was provided without the measurement of the pH and the description of packaging after storage.

However, based on the composition of the product and the results of the storage stability study, the product can be considered stable after 12 months at 30°C. Nevertheless, for confirmation and according to Guidance on the BPR, a new stability storage study in commercial packaging including appearance of the product, aspect of the packaging, pH, and AS content with a validated method of quantification should be provided in post-authorisation within two years. A shelf life of 12 months can be granted.

As no accelerated stability study and no stability study in low temperature were provided, risk mitigation measures below are proposed:

- do not store at temperature above 30°C,

- Protect from frost.

Physico chemical dangers have been assessed. It is classified H272 Oxid. Liquid 2. It is not explosive, and is not autoflammable in the conditions of uses.

No data have been provided on corrosive properties of the product. However, based on the composition of the product, non-corrosive properties are expected. For confirmation, a test UN Test in Part III of the UN Manual of Tests and Criteria, section 37.4 for the property corrosive to metal should be provided in post-authorization within two years.

Efficacy

The assessment of the product HYDROGEN PEROXIDE SOLUTION concerns the efficacy of the product as a disinfectant.

In accordance with the submitted tests and the requirements of the Guidance on the Biocidal Products Regulation Volume II Efficacy - Assessment and Evaluation (Parts B+C) (2018), the product HYDROGEN PEROXIDE SOLUTION is efficient against bacteria, bacterial spores, yeasts, fungi and virus when used within the STERRAD® 200 GMP system, for the disinfection of instruments and materials during manufacturing process in industries (PT02), except uses covered by the medical device Regulation.

The authorization holder has to report any observed resistance incidents to the Competent Authorities (CA) or other appointed bodies involved in resistance management.

Human health risk

The risk for the professional user using the product HYDROGEN PEROXIDE PRODUCT within the STERRAD® 200 GMP Sterilization System, is acceptable when gloves are worn during handling of product and when the following RMM is followed: “The user has to handle a hydrogen peroxide gas detector. If the threshold of 1.25 mg / m3 is exceeded, the room has to be evacuated”. Moreover, considering the classification of the product, we propose to add the following RMM in the SPC: “Use only in a well-ventilated area” and “In case of inadequate ventilation, wear respiratory protection”.

The risk for general public in contact with the STERRAD® 200 GMP System is acceptable.

Regarding the use claimed, food or feed contamination is not expected.

Environmental risk

For the use of the product HYDROGEN PEROXIDE SOLUTION within the STERRAD® 200 GMP System, the estimation of concentrations in aquatic (including sediment and STP) and soil compartments as well as in the groundwaters was not considered relevant.

The following risk mitigation measure: “Use only in the STERRAD system without tubes or pipes connected to sewage systems” must be implemented.

# ASSESSMENT REPORT

## Summary of the product assessment

### Administrative information

#### Identifier of the product / product family

| Identifier | Country (if relevant) |
| --- | --- |
| Hydrogen Peroxide Solution |  |

#### Authorisation holder

| Name and address of the authorisation holder | Name | Advanced Sterilization Products, Division of Cilag GmbH International |
| --- | --- | --- |
|  | Address | Gubelstrasse 34, 6300 Zug Switzerland |

#### Manufacturer(s) of the products of the family

| Name of manufacturer | Advanced Sterilization Products (ASP), Division of Cilag GmbH International |
| --- | --- |
| Address of manufacturer | Gubelstrasse 34 CH-6300 Zug Zurich Switzerland |
| Location of manufacturing sites | 33 Technology Drive, Irvine CA 92618 California United States |

#### Manufacturer(s) of the active substance(s)

| Active substance | Hydrogen peroxide |
| --- | --- |
| Name of manufacturer | PeroxyChem Spain, S.L.U |
| Address of manufacturer | C/Afueras s/n, ES-50784 La Zaida, Zaragoza Spain |
| Location of manufacturing sites | PeroxyChem Bayport Plant - 12000 Bay Area Blvd, TX 77507 Pasadena United States |

### Product (family) composition and formulation

NB: the full composition of the product according to Annex III Title 1 should be provided in the confidential annex.

Does the product have the same identity and composition as the product evaluated in connection with the approval for listing of the active substance(s) on the Union list of approved active substances under Regulation No. 528/2012?

Yes

No 	x

#### Identity of the active substance

| Main constituent(s) |  |
| --- | --- |
| ISO name | Hydrogen peroxide |
| IUPAC or EC name | Hydrogen peroxide |
| EC number | 231-765-0 |
| CAS number | 7722-84-1 |
| Index number in Annex VI of CLP |  |
| Minimum purity / content | 99.5% |
| Structural formula |  |

#### Candidate(s) for substitution

The active substance Hydrogen peroxide contained in the biocidal product HYDROGEN PEROXIDE SOLUTION is not a candidate for substitution in accordance with Article 10 of Regulation (EU) No.528/2012

#### Qualitative and quantitative information on the composition of the biocidal product

| Common name | IUPAC name | Function | CAS number | EC number | Content (%) |
| --- | --- | --- | --- | --- | --- |
| Hydrogen peroxide (Technical-TC) | Hydrogen peroxide | Active substance | 7722-84-1 | 231-765-0 | 59.8 |

#### Qualitative and quantitative information on the composition of the biocidal product family

Not concerned

#### Information on technical equivalence

Not concerned

#### Information on the substance(s) of concern

No substance of concern is identified in the product HYDROGEN PEROXIDE SOLUTION

#### Assessment of endocrine disruption (ED) properties of the biocidal product family

According to our assessment, none of the formulants contained in the product are identified as endocrine disruptors.

Please see confidential annex.

#### Type of formulation

Any other liquid- AL

### Hazard and precautionary statements

Classification and labelling of the products of the family according to the Regulation (EC) 1272/2008

| Classification |  |
| --- | --- |
| Hazard category | Ox. Liquid. 2 Acut tox.4 H302 Acute tox.2 H330 Skin Corr. 1B H314 Aquatic chronic cat 3 |
| Hazard statement | H272 May intensify fire; oxidizer H302: Harmful if swallowed H330: Fatal if inhaled H314: Causes severe skin burns and eye damage H412: Harmful to aquatic life with long-lasting effects |
| Labelling |  |
| Signal words | Danger |
| Hazard statements | H272 May intensify fire; oxidizer H302: Harmful if swallowed H330: Fatal if inhaled H314: Causes severe skin burns and eye damage H412: Harmful to aquatic life with long-lasting effects |
| Precautionary statements | P210 : Keep away from heat, hot surfaces, sparks, open flames and other ignition sources. – No smoking P220 : Keep/Store away from clothing/…/combustible materials. P260 : Do not breathe dust/fumes/gas/mist/vapours/spray P264 : Wash … thoroughly after handling. P270 : Do not eat, drink or smoke when using this product. P271 Use only outdoors or in a well-ventilated area. P280: Wear protective gloves/protective clothing/eye protection/face protection. P284 : [In case of inadequate ventilation] wear respiratory protection. P301 + P312: IF SWALLOWED: Call a POISON CENTER/doctor/…/if you feel unwell. P301 + P330 + P331: IF SWALLOWED: Rinse mouth. Do NOT induce vomiting. P303 + P361 + P353 : IF ON SKIN (or hair): Take off immediately all contaminated clothing. Rinse skin with water/ shower. P304 + P340 : IF INHALED : Remove person to fresh air and keep comfortable for breathing. P305 + P351 + P338 : IF IN EYES: Rinse cautiously with water for several minutes. Remove contact lenses if present and easy to do – continue rinsing. P310 : Immediately call a POISON CENTER/doctor/… P320 : Specific treatment is urgent (see … on this label) P321 : Specific treatment (see … on this label). P330 : Rinse mouth. P363 : Wash contaminated clothing before reuse. P370 + P378 : In case of fire: Use dry sand, dry chemical or alcohol-resistant foam for extinction. P403 + P233 : Store in a well ventilated place. Keep container tightly closed. P405 : Store locked up. P501 : Dispose of contents/container in accordance with local/regional/national/international regulation (to be specified)]. |
| Note | Supplemental hazard information: EUH071-“Corrosive to the respiratory tract” |

### Authorised use(s)

#### Use description

Table 1. Use # 1 – Professional user only in STERRAD® 200 GMP Systems

| Product Type | PT02 - Disinfectants and algaecides not intended for direct application to humans or animals (Disinfectants) |
| --- | --- |
| Where relevant, an exact description of the authorised use | The hydrogen peroxide solution will be used within the Sterrad® 200 GMP System for the disinfection of instruments and materials in manufacturing and other industrial sectors. |
| Target organism (including development stage) | Bacteria, bacterial spores, viruses, yeasts and fungi. |
| Field of use | Indoor |
| Application method(s) | Closed System To be applied only with the STERRAD® 200 GMP system. |
| Application rate(s) and frequency | Injection volume per injection stage: 5 mL of product HYDROGEN PEROXIDE SOLUTION Chamber concentration: 13.47 mg H2O2 /L. A disinfection cycle will have a minimum of 2 injection stages. The minimum volume of hydrogen peroxide used per cycle is 10 mL. |
| Category(ies) of users | Professional |
| Pack sizes and packaging material | Bottle volume/material: 500 mL HDPE |

#### Use-specific instructions for use

-

#### Use-specific risk mitigation measures

-

#### Where specific to the use, the particulars of likely direct or indirect effects, first aid instructions and emergency measures to protect the environment

-

#### Where specific to the use, the instructions for safe disposal of the product and its packaging

-

#### Where specific to the use, the conditions of storage and shelf-life of the product under normal conditions of storage

-

### General directions for use

#### Instructions for use

- Always read the label or leaflet before use and follow all the instructions provided. - Respect the conditions of use of the product (concentration, contact time, temperature, pH, etc.) - Clean and dry carefully all items before placing them in the STERRAD 200 GMP system. - Some materials and instruments are not compatible with hydrogen peroxide disinfection. Please refer to the STERRAD system Operator’s Manual for more information. - Inform the registration holder if the treatment is ineffective. - It is recommended to open packaging over a sink to ensure safety. - Before opening the outer packaging, check the product packaging for excessive damage during shipping. Check the tilt indicator on the side of the pail. Presence of blue beads inside the arrow indicates that the pail may have been mishandled or tipped. In either case, the inner contents may still be usable. Exercise caution when opening pail. Remove the cable tie that secures the cover to the pail. - To remove the cover, push back the top of the red lever with thumb while turning the cover counterclockwise. Carefully remove each bagged bottle from the pail and inspect to determine if any hydrogen peroxide has spilled into the yellow absorbent material. - Remove the cable tie wrapped around the outer bag. Keeping the package in an upright orientation, carefully remove the bottle and absorbent material from the bags. Remove the bottle from the absorbent material and inspect for any damage. - Remove the red protective cap from the bottle and inspect the white membrane inside the cap for damage. Note: The membrane is meant to vent vapour. Some condensation may be present but does not constitute a defect. - The white bottle cap is not intended to be removed. If contents are not damaged, replace the red protective cap and store the bottle in the original packaging if not for immediate use. If inner contents have spilled, refer to instructions for disposal and call local ASP Representative.

#### Risk mitigation measures

- Wear protective chemical resistant gloves (glove material to be specified by the authorisation holder within the product information), during product handling phase. - The user has to handle a hydrogen peroxide gas detector. If the threshold of 1.25 mg / m3 is exceeded, the room has to be evacuated. - Use only in a well-ventilated area. - In case of inadequate ventilation, wear respiratory protection. - Use only in the STERRAD system without tubes or pipes connecting to sewage systems

#### Particulars of likely direct or indirect effects, first aid instructions and emergency measures to protect the environment

- Inhalation: Remove victim to fresh air and keep at rest in a half-sitting position. Seek medical advice immediately if symptoms occur and/or large quantities have been inhaled. - Skin contact: Remove contaminated clothing and shoes. Wash contaminated skin with water. Contact poison treatment specialist if symptoms occur. - Eye contact: Immediately flush with plenty of water, occasionally lifting the upper and lower eyelids. Check for and remove any contact lenses if easy to do. Continue to rinse with warm water for at least 10 minutes. Get medical attention if irritation or vision impairment occurs. - Mouth contact: Wash out mouth with water. Contact poison treatment specialist immediately if symptoms occur and/or in case of mouth contact with large quantities. - Do not give fluids or induce vomiting in case of impaired consciousness; place in recovery position and seek medical advice immediately. - Keep the container or label available.

#### Instructions for safe disposal of the product and its packaging

- Do not discharge unused product on the ground, into water courses, into pipes (sink, toilets…) nor down the drains. - Dispose of unused product, its packaging and all other waste, in accordance with local regulations.

#### Conditions of storage and shelf-life of the product under normal conditions of storage

- Protect from frost. - Do not store at temperature above 30°C. - Shelf life: 1 year.

### Other information

- The authorization holder should report any observed incidents related to the efficacy to the Competent Authorities (CA). - The use for the disinfection of instruments and materials which is covered by the medical devices Regulation is not concerned by this authorisation.

### Packaging of the biocidal product

| Type of packaging | Size/volume of the packaging | Material of the packaging | Type and material of closure(s) | Intended user (e.g. professional, non-professional) | Compatibility of the product with the proposed packaging materials (Yes/No) |
| --- | --- | --- | --- | --- | --- |
| Bottle | 0.5 L | HDPE | / | Professional | Yes |

### Documentation

#### Data submitted in relation to product application

Physico-chemical properties,

Physico-chemical properties studies and analytical methods on the biocidal product Hydrogen peroxide solution were provided by Advanced Sterilization Products.

Efficacy

New data on the product is submitted for the demonstration of the efficacy. Please refer to the list of reference.

Human health and Environmental risk

No new data have been submitted for the product HYDROGEN PEROXIDE SOLUTION.

#### Access to documentation

The applicant has submitted a letter of access signed by PeroxyChem a member of the hydrogen peroxide task force, which represents the owners of the studies and the dossier on hydrogen peroxide that were submitted in support of the approval of  hydrogen peroxide as a biocidal active substance under the Biocidal Products Regulation..

## Assessment of the biocidal product

### Intended use(s) as applied for by the applicant

Table 1. Intended use # 1 – name of the use

| Product Type | PT02 - Disinfectants and algaecides not intended for direct application to humans or animals (Disinfectants) |
| --- | --- |
| Where relevant, an exact description of the authorised use | The proposed field of use of hydrogen peroxide solution is the control of bacteria, bacterial spores, yeasts, fungi and viruses in the industrial / manufacturing sector. The product is applied by vaporising hydrogen peroxide in the STERRAD® 200 GMP Sterilisation System, a low temperature sterilizer using the STERRAD process. The proposed uses are included in PT02. The biocidal product contains hydrogen peroxide as active substance which is antimicrobial active against a wide range of microorganisms: bacteria, fungi, moulds, viruses and spores. The efficacy of hydrogen peroxide results from its oxidising property: the hydroxyl radical attacks the lipids of the membrane, the DNA or other components essential to the cell. The product is intended for professional use only. |
| Target organism (including development stage) | Broad spectrum microbiocidal activity against bacteria, viruses, and fungi. - Bacillus spp.-Bacterial spores¦Bacteria-Bacteria - Bacillus spp.-Bacterial spores¦Bacteria-Bacteria, aerobic Gram-negative - Aerobacter spp.-Bacterial spores¦Bacteria-Bacteria, aerobic Gram-negative - Pseudomonas aeruginosa-Bacterial spores¦Bacteria-Bacteria, aerobic Gram-negative - Escherichia coli-Bacterial spores¦Bacteria-Bacteria, aerobic Gram-negative - Serratia marcescens- - Moroxelia osloensis-Spores and spore producing structures-Bacteria, aerobic Gramnegative - Mycobacterium bovis-Spores and spore producing structures-acid - Candida albicans-Spores and spore producing structures-Yeast - Candida parapsilosis-Spores and spore producing structures-Yeast - Trichophyton mentagrophytes-Spores and spore producing structures-filamentous algae - Aspergillus niger-Spores and spore producing structures-Filamentous fungi |
| Field of use | Indoor |
| Application method(s) | Closed System The Sterrad® process is a plasma system that uses hydrogen peroxide as the source of active species. The process, which can be described in five phases, uses a gas-diffusion phase to penetrate all parts of the load before the plasma is created. The five phases are: - Vacuum phase: The chamber is evacuated, reducing internal pressure in preparation for the subsequent reaction; - Injection phase: A measured amount of liquid peroxide is injected into the chamber, evaporating the aqueous hydrogen peroxide solution and dispersing it into the chamber, where it kills bacteria on any surface it can reach; - Diffusion phase: The hydrogen peroxide vapour permeates the chamber, exposing all load surfaces to the sterilant and rapidly sterilizes devices and materials without leaving any toxic residues. At the completion of this phase, the chamber pressure is reduced, and the plasma discharge is initiated; - Plasma phase: An electromagnetic field is created in which the hydrogen peroxide vapour breaks apart, producing a low temperature plasma cloud that contains ultraviolet light and free radicals. Following the reaction, the activated components lose their high energy and recombine to form oxygen and water; - Vent phase: The chamber is vented to equalize the pressure enabling the chamber door to be opened. There is no need for aeration or cool-down. Devices are ready for immediate use. Phases 1, 2 and 3 are then run for a second time for added efficacy, reprocessing to assure optimal sterilization. The plasma phase seeks to reduce residual hydrogen peroxide on the process load. This step degrades peroxide to oxygen and water and reduces residuals on the product being sterilized. When a plasma contacts the surface of an item to be sterilized, the collisions between the active species and other molecules cause a significant proportion of the active species to return to the ground state and recombine to form oxygen and water. |
| Application rate(s) and frequency | Typical volume injected into the chamber per injection stage is approximately 5 mL which is the pre-defined volume of the metering system. A sterilisation cycle will have a minimum of 2 injection stages. The minimum volume of hydrogen peroxide used per cycle is 10 mL. - The in-use product contains 59% w/w hydrogen peroxide and 41% w/w water. The nominal container (bottle) capacity is 0.5 L. The nominal fill mass per container is 546 ± 22 g solution with a target volume of 444 ± 18 mL solution (1.23 g/mL solution density at 23°C) on the date of manufacture - The typical volume of each injection into the STERRAD system is about 5 mL which is the pre-defined volume of the metering system. A typical cycle will have a minimum of 2 injection stages or more per cycle. Therefore, the minimum volume of H2O2 used per cycle will be about 10 ml. |
| Category(ies) of users | Professional |
| Pack sizes and packaging material | Bottle volume/material: 0.5 L HDPE Nominal bottle wall thickness: 1.52 mm Screw closure size: 53mm Screw closure material: Polypropylene (PP) Cap: LDPE |

### Identity

The biocidal product is not the same as the one assessed for the inclusion of the active substances in annex 1 of directive 98/8/EC. The composition of the product is presented in a confidential annex. The product contains 59.8% of technical active substance hydrogen peroxide (59.5% of pure active substance hydrogen peroxide).

The product does not contain PT6 preservative.

The product is not diluted for use.

Formulation type: Any other liquid (AL)

Hydrocarbon and H304 cat.1 co-formulant content: <10%

### Physical, chemical and technical properties

| Property | Guideline and Method | Purity of the test substance (% (w/w) | Results | FR Evaluation | Reference |
| --- | --- | --- | --- | --- | --- |
| Physical state at 20 °C and 101.3 kPa |  | Not available | liquid | Acceptable | Raw material specification – Hydrogen Peroxide Solution – 59%. Report Number 37-15004-002 |
| Colour at 20 °C and 101.3 kPa |  | Not available | colourless | Acceptable | Raw material specification – Hydrogen Peroxide Solution – 59%. Report Number 37-15004-002 |
| Odour at 20 °C and 101.3 kPa |  | Not available | odourless | Acceptable | Raw material specification – Hydrogen Peroxide Solution – 59%. Report Number 37-15004-002 |
| Acidity / alkalinity |  | Not available | pH: 4.0 – 8.0 | Not Acceptable | Raw material specification – Hydrogen Peroxide Solution – 59%. Report Number 37-15004-002 |
|  | CIPAC MT 75.3 | H2O2 58.7% w/w ESRS 25/19 | At 20°C, pH of 0.1% w/v hydrogen peroxide preparation = 5.29 ± 0.10 At 20°C, pH of undiluted hydrogen peroxide test item = 0.89 ± 0.02 | Acceptable | Hydrogen Peroxide: Determination of Physicochemical Properties– Thomas D., 2019 Study Number 3202413-2019 |
| Relative density / bulk density |  | Not available | 1.2315 g/ml at 20°C | The relative density obtained by calculation is not acceptable. | Ref supplier provided calculator: http://h2o2.evonik.com/product/h2o2/en/about-hydrogen-peroxide/basic-information/calculations/pages/calculate.aspx. Ref: W. C. SCHUMB, C. N. SATTERFIELD. and R. L. WENTWORTH. Hydrogen peroxide. Reinhold Publishing Corp., New York. 1955. Chap. 5t |
|  | OECD 109 | H2O2 58.7% w/w ESRS 25/19 | 1.236 kg/m3 at 20°C | Acceptable | Hydrogen Peroxide: Determination of Physicochemical Properties– Thomas D., 2019 Study Number 3202413-2019 |
| Storage stability test – accelerated storage | / | / | No data provided. | Not acceptable As the accelerated storage stability study was not provided the mention « do not store at temperature above 30°C » is added to mitigation measure in SPC. | / |
| Storage stability test – long term storage at ambient temperature |  | Hydrogen Peroxide for 200GMP Sterilizer Lot 217141 Analitycal method for AS content is titrimetric method (TP-25001-001) or Densitometric Method (TP-25191-001). | Storage at 30°C “In Shipping Pail” is defined as 2 bottles within a Pail, the final package configuration. The “Out of Shipping Pail” is defined as an individual bottle. The material of the bottle is not precised. Time (months); T0; T3; T6; T9; T11; T12; T13. AS content in the shipping pail %; 58.8; 58.5; 58.2; 58; 57.7; 57.8; 57.5. Variation of AS (%); /; -0.5; -1; -1.4; -1.9; -1.7; -2.2. AS content out of the shipping pail %; 58.8; 58.2; 57.7; 57.2; 56.6; 56.1; 55.7. Variation of AS (%); /; -1; -1.9; -2.7; -3.8; -4.6; -5.3. pH. | The pH should have been measured and the aspect of the product and the packaging before and after storage should have been described. The packaging tested in the storage stability is not reported and should have been precised. Nevertheless, based on the composition of the product, and the result of the storage stability study and the pH at T0, storage duration of 1 year is considered acceptable. However according to Guidance on the BPR and for confirmation a new stability storage study in commercial packaging including appearance of the product, aspect of the packaging, pH, and AS content with a validated method of quantification should be provided in post-authorisation within two years. | AUTEUR (2015), study number R-17027 |
| Storage stability test – low temperature stability test for liquids | / | / | No data provided. | As study of storage stability at low temperature was not provided, the mention « protect from frost » must be added in SPC. | / |
| Effects on content of the active substance and technical characteristics of the biocidal product - light | / | / | Not provided | Not relevant, absorption spectrum shows no propensity of the molecule to be decomposed by UV/Vis light. | / |
| Effects on content of the active substance and technical characteristics of the biocidal product – temperature and humidity | / | / | Not provided | Not relevant, there are mitigation measures for storage conditions. | / |
| Effects on content of the active substance and technical characteristics of the biocidal product - reactivity towards container material | / | / | Not provided | The aspect of packaging during shelf life should be provided in post-authorization. | / |
| Wettability | / | / | / | Not relevant for an AL formulation | / |
| Suspensibility, spontaneity and dispersion stability | / | / | / | Not relevant for an AL formulation | / |
| Wet sieve analysis and dry sieve test | / | / | / | Not relevant for an AL formulation | / |
| Emulsifiability, re-emulsifiability and emulsion stability | / | / | / | Not relevant for an AL formulation | / |
| Disintegration time | / | / | / | Not relevant for an AL formulation | / |
| Particle size distribution, content of dust/fines, attrition, friability | / | / | / | Not relevant for an AL formulation | / |
| Persistent foaming | / | / | / | Not relevant for an AL formulation | / |
| Flowability/Pourability/Dustability | / | / | / | Not relevant for an AL formulation | / |
| Burning rate — smoke generators | / | / | / | Not relevant for an AL formulation | / |
| Burning completeness — smoke generators | / | / | / | Not relevant for an AL formulation | / |
| Composition of smoke — smoke generators | / | / | / | Not relevant for an AL formulation | / |
| Spraying pattern — aerosols | / | / | / | Not relevant for an AL formulation | / |
| Physical compatibility | / | / | / | Not relevant for an AL formulation | / |
| Chemical compatibility | / | / | / | Not relevant for an AL formulation | / |
| Degree of dissolution and dilution stability | / | / | / | Not relevant for an AL formulation | / |
| Surface tension | / | / | ~60 dyn/cm at 20°C | Not acceptable Technical characteristic from MSDS and no measure has been performed. The data are not accurate enough and are necessary for the control. | M.K. Phibbs and P.A. Giguere. Can. J. Chem., 29:173 (1951): http://www.nrcresearchpress.com/doi/pdfplus/10.1139/v51-022 |
|  | OECD 115 | H2O2 58.7% w/w ESRS 25/19 | 75.25 mN/m | Acceptable | Hydrogen Peroxide: Determination of Physicochemical Properties– Thomas D., 2019 Study Number 3202413-2019 |
| Viscosity | / | / | ~1.2 cP | Not acceptable Technical characteristic from MSDS and no measure has been performed. The data are not accurate enough and are necessary for the control. | http://www.h2o2.com/technical-library/physical-chemical-properties/physical-properties/default.aspx?pid=16&name=Viscosity-of-H2O2-Solutions |
|  | OECD 114 | H2O2 58.7% w/w ESRS 25/19 | Kinematic viscosity at 20°C: 0.9823 mm²/s Kinematic viscosity at 40°C: 0.6934 mm²/s Dynamic viscosity at 20°C: 1.2141 mPa.s Dynamic viscosity at 40°C: 0.8570 mPa.s | Acceptable | Hydrogen Peroxide: Determination of Physicochemical Properties– Thomas D., 2019 Study Number 3202413-2019 |

Conclusion on the physical, chemical and technical properties of the product

The product HYDROGEN PEROXIDE SOLUTION is an all other liquid (AL) formulation. The stability data indicate a shelf life of at least 1 year at ambient temperature. However, the pH should have been measured and the aspect of the product and the packaging before and after storage should have been described. Additionally, the packaging tested in the storage stability is not reported and should have been precised. However, based on the composition of the product, storage duration of 1 year is considered acceptable. Nevertheless, according to Guidance on the BPR and for confirmation, a new stability storage study in commercial packaging including appearance of the product, aspect of the packaging, pH, and AS content with a validated method of quantification should be provided in post-authorisation within two years. Mitigation measure: Protect from frost. Do not store at temperature above 30°C.

### Physical hazards and respective characteristics

| Property | Guideline and Method | Purity of the test substance (% (w/w) | Results | FR evaluation | Reference |
| --- | --- | --- | --- | --- | --- |
| Explosives | Statement | / | According to the classification of all present ingredients in this biocidal product family under the Regulation 1272/2008 (CLP), which do not classify as Explosives, it is not necessary to determine the explosive properties of any of the biocidal products within the family. Regarding H2O2 the predominant hazard is the oxidising property | Acceptable | / |
| Flammable gases | / | / | Not applicable it is a liquid. | / | / |
| Flammable aerosols | / | / | Not applicable it is a liquid. | / | / |
| Oxidising gases | / | / | Not applicable it is a liquid. | / | / |
| Gases under pressure | / | / | Not applicable it is a liquid. | / | / |
| Flammable liquids | Statement | / | Inorganic oxidising liquids are not flammable and therefore do not have to be subjected to the classification procedures for the hazard classes flammable liquids or pyrophoric liquids | Acceptable | / |
| Flammable solids | / | / | Not applicable it is a liquid. | / | / |
| Self-reactive substances and mixtures | / | / | Not applicable it is a liquid. | / | / |
| Pyrophoric liquids | Statement | / | Inorganic oxidising liquids are not flammable and therefore do not have to be subjected to the classification procedures for the hazard classes flammable liquids or pyrophoric liquids | Acceptable | / |
| Pyrophoric solids | / | / | Not applicable it is a liquid. | / | / |
| Self-heating substances and mixtures | / | / | Not applicable it is a liquid. | / | / |
| Substances and mixtures which in contact with water emit flammable gases | / | / | Not applicable it is a liquid. | / | / |
| Oxidising liquids | Statement | / | Biocide product contains 59% H2O2, according to harmonized classification of H2O2, this biocide product should be classified : Ox. Liq. 2; H272:50 % ≤ C < 70 % | Acceptable Biocidal product is classified as Liq.2 H272 | / |
| Oxidising solids | / | / | Not applicable it is a liquid. | / | / |
| Organic peroxides | / | / | Not applicable it is a liquid. | / | / |
| Corrosive to metals | / | / | No data provided. | The biocidal product does not contain halogen, able to form complexes with metals, is not a strong acid. Therefore, based on the composition the product could be considered as non-corrosive. However, this should be confirmed. A test UN Test in Part III of the UN Manual of Tests and Criteria, section 37.4 for the property corrosive to metal should be provided in post-authorization within two years. | / |
| Auto-ignition temperatures of products (liquids and gases) | / | / | No data provided. | Acceptable | / |
| Relative self-ignition temperature for solids | / | / | Not applicable it is a liquid. |  | / |
| Dust explosion hazard | / | / | Not applicable it is a liquid. |  | / |

Conclusion on the physical hazards and respective characteristics of the product

The product is neither flammable nor auto-flammable. It has no explosive properties but has oxidising properties. No data was provided for corrosive property. Based on the composition of the product, non-corrosive properties are expected. Nevertheless, for confirmation a test UN Test in Part III of the UN Manual of Tests and Criteria, section 37.4 for the property corrosive to metal should be provided in post-authorization within two years. Classification of the product related to physico-chemical properties : Ox. Liq. 2; H272: May intensify fire; oxidiser

### Methods for detection and identification

Report: Supplemental Analytical Method Validation – Test Procedures TP-25001-001, TP-25091-001, and TP-25108-001 – All Procedures using the Hydrogen Peroxide Concentration Solutions: Titrametric Method   2014

Report no R-15067 + appendices

Test facilities: not reported by applicant

Principle of the method:

Hydrogen peroxide and an excess of potassium iodide react in an acidic solution to generate iodine. The liberated iodine is titrated with standard sodium thiosulfate using starch as the indicator. The concentration of the hydrogen peroxide solution can be determined from the weight of sample titrated and the volume of sodium thiosulfate required to reach the titration endpoint.

Validation data:

| Specificity | To demonstrate the specificity of the method, several solutions are analyzed: - perchloric acid contaminated samples - deionized water only samples - 2-propanol contaminated samples - Test item of the product Variations in standard deviation, RSD, and concentration were determined. For hydrogen peroxide percent concentration, the average of the five titrations is 59.11%. The % RSD is 0.1%. Results are within the acceptable criteria for specificity. |  |
| --- | --- | --- |
| Linearity | Linearity was studied by carrying out twelve concentrations between 45% to 65% of H2O2. Calibration curve has been provided with a R2 higher than 0.99. |  |
|  | Compound | Linearity % |
|  | Active substance | 45 to 65% Y = 58624 x – 0.0679 R2 = 0.9998 n=1 |
| Precision | Repeatability was evaluated by analyzing fifth twelve test item solution (45 to 65% of H2O2) but only one level of concentration is presented here. |  |
|  | Compound | Repeatability (RSD) |
|  | Active substance H2O2 (59.5%) | RSD = 0.184% |
| Accuracy | Accuracy was determined by analysis of 5 reconstituted samples. The accuracy results are expressed as the recovery rate. Twelve levels of concentration of test item solution (45 to 65% of H2O2) are performed but only three levels of concentration are presented here. Fortification level; Recovery range; Mean recovery rate; RSD (%); n. 58%; 99.56-9.98; 99.8; 0.2; 5. 59.5%; 99.63-100.1; 99.8; 0.18; 5. 62.5%; 99.87-100.07; 100; 0.07; 5. |

The analytical method is fully validated for the determination of the active substance H2O2 in the product.

Another method to quantify AS has been provided by applicant and evaluated below:

Report: Supplemental Analytical Method Validation – Test Procedure TP-25191-001 – Hydrogen Peroxide Concentration of Solutions, Densitometric Method 2014

Report no R-15068 + appendices

Test facilities: not reported by applicant

Principle of the method:

A densitometer is used to measure the density of solutions containing the active substance in the biocidal product using the oscillating body method, which derives density measurements based on the frequency of oscillation of the testing sample in the measuring device.

Validation data:

| Specificity | To demonstrate the specificity of the method, several solutions are analyzed: - perchloric acid contaminated samples - deionized water only samples - 2-propanol contaminated samples - Test item of the product Variations in standard deviation, RSD, and concentration were determined. For hydrogen peroxide percent concentration, the average of the five titrations is 59.11%. The % RSD is 0.1%. Results are within the acceptable criteria for specificity. |  |
| --- | --- | --- |
| Linearity | Linearity was studied by carrying out eighteen concentrations between 45% to 70% of H2O2. Calibration curve has been provided with a R2 higher than 0.99. |  |
|  | Compound | Linearity % |
|  | Active substance | 45 to 70% Y =0.9684x + 1.9869 R2 = 0.9998 n=1 |
| Precision | Repeatability was evaluated by analyzing trice eighteen test item solution (45 to 65% of H2O2) but only three levels of concentration is presented here. |  |
|  | Compound | Repeatability (RSD) |
|  | Active substance H2O2 (59%) | RSD = 0.03% |
|  | Active substance H2O2 (59.5%) | RSD = 0.04% |
|  | Active substance H2O2 (60%) | RSD = 0.03% |
| Accuracy | Accuracy was determined by analysis of 3 reconstituted samples. The accuracy results are expressed as the recovery rate. Eighteen levels of concentration of test item solution (45 to 70% of H2O2) are performed but only four levels of concentration are presented here. Fortification level; Recovery rate; Mean recovery rate; RSD (%); n. 58%; 100.32; 100.35; 100.29; 100.32; 0.03; 3. 59%; 100,34 ;100,3; 100,35; 100.3; 0.03; 3. 59.5%; 99,4; 99,45; 99,47; 99.4; 0.04; 3. 60%; 100,23; 100,26;100,21; 100.2; 0.03; 3. |

The analytical method is fully validated for the determination of the active substance H2O2 in the product.

Analytical methods for H2O2 residues in soil, air, water (drinking water) and sediment are available in Assessment Report of H2O2 Product-type 2, Feb. 2017. The applicant Advanced has a Letter of Access from PeroxyChem Spain s.l.u. for these data.

Based on the intended uses of the product, no contamination of the environment is foreseen. Analytical methods for H2O2 residues in soil, air, water (including drinking water) and sediment are unnecessary.

As the active substance H2O2 is not classified Toxic or Very Toxic, an analytical method for the determination of hydrogen peroxide residue in human body fluids and tissues is unnecessary.

No analytical method in soil is required according to the Assessment report of hydrogen peroxide.

As the product hydrogen peroxide solution is not intended to be used with surface in contact with food/feed of plant and animal origin, analytical method for the determination of hydrogen peroxide residue in food/feed of plant and animal origin is unnecessary.

Conclusion on the methods for detection and identification of the product

The analytical method is fully validated for the determination of the active substance hydrogen peroxide in the product. Analytical methods were provided at EU level for the determination of active substance residue in soil, water and air with respectively LOQ =0.14 mg/m3 and 0.01 µg/L. Hydrogen peroxide is not toxic (T) or very toxic (T+) active substance. Therefore, an analytical method in biological matrices is not required. No analytical method in soil is required according to the Assessment report of hydrogen peroxide. The product is not intended to be used on surface in contact with food/feed of plant and animal origin, analytical method for the determination of hydrogen peroxide in food/feed of plant and animal origin is not required.

### Efficacy against target organisms

#### Function and field of use

MG 01: Disinfectants

PT2: Disinfectants and algaecides not intended for direct application to humans or animals

Hydrogen peroxide solution is intended to be used within the STERRAD® 200 GMP System with the final purpose of sterilizing instruments and materials during their manufacturing in the industrial sector.

It has to be noted that in the frame of application for marketing authorisation of HYDROGEN PEROXIDE SOLUTION as a biocidal product, only the use as a disinfectant has been considered. Moreover, the use for the disinfection of instruments and materials which is covered by the medical devices Regulation is not concerned.

The product is used by professional users in industry.

#### Organisms to be controlled and products, organisms or objects to be protected

The claimed spectrum of activity includes bacteria, bacterial spores, viruses, fungi and yeasts on instruments and materials during their manufacturing in the industrial sector.

The product is used for the purpose of the protection of human health.

According to the applicant, the product can be used on a variety of articles in an industrial setting.

The customers range from medical device companies to specialty packaging manufacturers. The articles processed may range from specialty packaging material to surgical implants.

#### Effects on target organisms, including unacceptable suffering

The product is able to produce a reduction in the number of viable bacterial cells (bactericidal activity), of viable bacterial endospores (sporicidal activity), of yeast cells (yeasticidal activity), of moulds spores (fungicidal activity) and of infectious virus particles (virucidal activity) of relevant test organisms under defined conditions.

#### Mode of action, including time delay

Hydrogen peroxide is reactive, and it degrades rapidly in contact with organic material. A significant proportion of hydrogen peroxide decomposes to water and oxygen.

The antimicrobial action of hydrogen peroxide stems from its ability to form powerful oxidants such as the hydroxyl radical and singlet oxygen. These reactive oxygen species cause irreversible damage to a host of cell components such as enzymes, membrane constituents and DNA.

#### Efficacy data

The applicant has provided following justifications relative to the devices, application rate and test organisms used in the submitted efficacy studies:

Equipment

Advanced Sterilization Products (ASP) has manufactured a range of STERRAD® sterilizers intended for healthcare facilities, scientific applications and industrial sterilization. Sizes have ranged from approximately 50 L to 200 L chambers. Each model sterilizer provides similar features and all run sterilization processes composed of similar “building block” process steps. All sterilizers are capable of delivering a 10-6 sterility assurance level (SAL) utilizing similar sterilization cycles.

For the uses considered in this dossier, only one model is claimed, STERRAD® 200 GMP. The STERRAD® 200 GMP have programmable sterilization cycles which are developed and validated by the end user.

The STERRAD® 200 GMP is designed based on the STERRAD® 200, which, in turn, was designed based existing process knowledge derived from a previous model in the STERRAD® family of products, the STERRAD 100S. Data derived from the previous models are used to support the efficacy of the H2O2 in the process.

FR CA accept applicant’s justification

Description of the process

All STERRAD® cycles, regardless of chamber volume have a number of common aspects in the available sterilization process steps.

The basic process steps are:

- Initial Vacuum – Air is removed from the sterilization chamber

- Air Plasma – A plasma is lit (no peroxide)

- Vent – Filtered air enters the chamber

- Vent - The sterilizer is held at atmospheric pressure

- Vacuum – Air is removed from the sterilization chamber

- Injection (sometimes called Transfer) – Hydrogen peroxide solution is vaporized into the sterilization chamber and the materials/instruments exposed for a period of time. This step results in destruction of microorganisms.

- Diffusion – Filtered air enters the chamber and the sterilizer is held at atmospheric pressure. This step provides additional efficacy against any surviving microorganisms.

- Vacuum – Air and hydrogen peroxide vapour are evacuated from the chamber

- Plasma (with peroxide) – A plasma is lit with air and hydrogen peroxide. This step degrades peroxide to oxygen and water and reduces H2O2 residuals on the product being sterilized.

Note: Steps E-I are considered as a “block”. The number of blocks per cycle type varies by Sterrad® model, and for industrial applications (Model 200 GMP) the owner defines the process steps.

The hydrogen peroxide solution is delivered within the STERRAD® 200 GMP from a peroxide container held inside the system.

The Air Plasma step (B), if used, is intended to warm the load. Heat is generated by the air plasma which in turn is absorbed by the load and will raise the load temperature a few degrees.

Steps F&G constitute the hydrogen peroxide exposure steps. Complete inactivation of microorganisms occurs in these steps.

Step H is an evacuation step to remove most hydrogen peroxide vapour and filtered air admitted during Step G.

Step I is the plasma step with the remaining hydrogen peroxide vapour not removed by step H. Step I is to reduce residual hydrogen peroxide in the chamber and on the load.

A final process step, P (venting), is required for all sterilizers to bring the system to atmospheric pressure and allow the door to open.

A generalized schematic of a two block STERRAD® cycle is presented below to help visualize the process.

Indicator Organism

Geobacillus stearothermophilus has been determined to be the most resistant organism to sterilization with the STERRAD® Hydrogen Peroxide Gaz Plasma Sterilization System,,. This was demonstrated during the early development of the STERRAD® System. Several independent researchers have independently demonstrated the resistance of G. stearothermophilus to hydrogen peroxide-based sterilization. Block’s textbook on sterilization and disinfection lists G. stearothermophilus as the most resistant organism to sterilization by vapour phase hydrogen peroxide. The study by Kokubo et al. demonstrated that environmental isolates of several Bacillus species were much less resistant to sterilization by vapor phase hydrogen peroxide than commercially prepared B. stearothermophilus spores. The European Pharmacopoeia and the FDA lists B. stearothermophilus as the most resistant organism to sterilization by hydrogen peroxide.

Therefore, efficacy studies only on Geobacillus stearothermophilus have been accepted to cover other targets organisms claimed (bacteria including spores, yeasts, fungi and viruses).

Application rate

The minimum injection volume of H2O2 per cycle claimed is 10 mL (a minimum of 2 injections at 5 mL per phase).

The nominal target concentration in the chamber can be calculated as follows:

Concentration in the chamber (mg/L) = (weight of H2O2 delivered (gr) x 1000 x conc. in %)/ total volume of the chamber in litres.

For the STERRAD® 200 GMP:

- 5 ml of H2O2 injected at 59 % w/w weight 6.09 g.

- The total volume of the chamber is 268 L.

Chamber Concentration calculation:

Concentration chamber (mg/L) = (6.09 x 1000 x 59%) / 268 = 13.41 mg H2O2/L

In practice, the amount of H2O2 injected may be greater than presented in the example. When incorporated into customer manufacturing process, the loads cycles are independently designed and validated at the customer site. These customized cycles are owned by the individual customers and the applicant does not have access to the data.

| Function | Field of use envisaged | Test substance | Test organism(s) | Test method | Test system / concentrations applied / exposure time | Test results: effects | Reference |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Disinfection | Instruments and materials in hospitals and industry (PT2) | Hydrogen peroxide (59% nominal H2O2) (within the STERRAD® 100, total chamber volume: 173L) | HIV-1 virus (strain HTLV-IIIB) Human Hepatitis A Virus (strain HM-175, variant 18F of VR-1073) Respiratory Syncytial Virus (strain Long, VR-26) Vaccinia virus (strain WR, VR-119) Herpes Simplex Virus (strain F-1, VR733) Poliovirus Type 2 (strain Lansing VR-1002) | Films of each virus were prepared by spreading 0.2 mL viral suspension on the bottom of 100 x 150 mm sterile glass Petri dishes. The virus films were left to dry at room temperature prior to the Petri dishes being placed into trays and the trays double-wrapped in a heavy-duty polypropylene wrap. | Two 25-minute half-cycles, each with an injection volume of 1.8 mL of 59% hydrogen peroxide followed by 7.5 minutes of plasma at 400 W. 4 replicates of each virus were tested. Hydrogen peroxide concentration in the chamber: 1.8 x 1.23 = 2.2140 g (2.2140 x 1000 x 59%)/173 = 7.55 mg/L | Results of hydrogen peroxide gas plasma system against test viruses. Effect on median tissue culture infective dose (TCID50) values: Input control; Dried virus control; Cytotoxicity control; Test samples. HIV-1. Half cycle 1; 106.5; 106.5; 102.5; 102.5. Half cycle 2; 106.5; 106.25; 102.5; 102.5. HAV. Half cycle 1; 104.5; 104.5; 100.5; 100.5. Half cycle 2; 104.5; 104.5; 100.5; 100.5. RSV. Half cycle 1; 105.0; 104.0; 101.5; 101.5. Half cycle 2; 104.5; 103.75; 100.5; 100.5. Vaccinia. Half cycle 1; 105.0; 105.0; 100.5; 100.5. Half cycle 2; 105.0; 104.5; 100.5; 100.5. HSV-1. Half cycle 1; 107.0; 105.5; 101.5; 101.5. Half cycle 2; 107.5; 105.25; 101.5; 101.5. Poliovirus type 2. Half cycle 1; 107.5; 105.5; 100.5; 100.5. Half cycle 2; 107.0; 105.5; 100.5; 100.5. Virucidal activity of the product, when used within the Sterrad 100 system, has been demonstrated against the tested organisms, at a chamber concentration of 7.55 mg/L of hydrogen peroxide (corresponding to an injection of 1.8 mL of 59% hydrogen peroxide per injection stage). | Roberts C. and Antonoplos P., 1998.* RI = 2 |
| Disinfection | Instruments and materials in hospitals and industry (PT2) | Hydrogen peroxide (59% nominal H2O2) (within the STERRAD® 100S, total chamber volume: 173L) | Geobacillus stearothermophilus ATCC 7953 | Death kinetics 5 replicates per spore carrier material at each hydrogen peroxide volume tested. The carriers were placed in open glass Petri dishes placed on top of instruments in each of two trays. The validation load was then double wrapped with Heavy Duty Spunguard. Recovery methods: - plate count for concentrations of 0.02, 0.045, 0.09 and 0.135 mL - direct culture for concentrations of 0.18, 0.27, 0.36, 0.72 and 1.44 mL. Controls: One positive control (unprocessed, inoculated spore carrier, for each material type) was incubated in TSB each day testing was conducted. One negative control (media) was incubated each day testing was conducted. Bacteriostasis testing was performed for each material. | One half-cycle, with an injection volume of 0.02 to 1.44 mL of 59% hydrogen peroxide. Preconditioning = 10 min preconditioning plasma Half-cycle parameters set to: Injection = 6 min Diffusion = 2 min Plasma = 2 min Spore carriers (approx. 5 mm x 20 mm) inoculated with at least 1 x 106 G. stearothermophilus spores. Spore carriers were of the following material types: PTFE (PT) Polyurethane (PU) Polyethylene (PE) Silicone (S) Glass filters (GF) Stainless steel (SS) Monel (M) Example of Hydrogen peroxide concentration in the chamber: 1.440 x 1.23 = 2.2140 g (2.2140 x 1000 x 59%)/173 = 6.04 mg/L | Mean Log10 values for B. stearothermophilus. Spore survivors with increasing H2O2 injection volumes. H2O2 injection Volume (µL); H2O2 conc. - chamber (mg/L); Material. SS; S; PU; M; GF; PT; PE. 20; 0.08; 5.65; 4.43; 4.53; 5.11; 4.67; 5.26; 5.04. 45; 0.19; 5.36; 5.04; 4.59; 4.82; 4.84; 5.26; 5.08. 90; 0.38; 4.83; 4.64; 4.46; 4.34; 4.34; 5.08; 4.82. 135; 0.57; 4.52; 4.08; 3.99; 4.54; 3.97; 4.23; 4.52. 180; 0.76; na; na; na; na; na; na; na. 270; 1.13; na; 0.21; na; na; na; na; na. 360; 1.51; na; na; na; na; na; na; na. 720; 3.02; na; na; na; na; na; na; na. 1440; 6.04; na; na; na; na; na; na; na. na – Log10 not calculated, no survivors for 720 and 1440 µL or 5/5 growth 180 to 360 µL, after 21 days Initial concentrations (log10 CFU per carrier): SS = 6.01, S = 6.15, PU = 6.08, M = 6.38, GF = 6.20, PT = 6.08, PE = 6.10 Controls: Positive controls demonstrated outgrowth of B. stearothermophilus and negative controls were sterile following incubation. No bacteriostatic effect from the process for any of the materials tested. Endospore death kinetics data obtained for each material tested demonstrates a positive “dose response” to increasing volumes of hydrogen peroxide injected at half cycle conditions. There were no spore survivors on any material with an injection volume of 720 µL or greater. The nominal Hydrogen peroxide volume for the half cycle is 1440 µL corresponding to an H2O2 concentration of 6.04 mg/L in the chamber. The results obtained in this test, at a minimum, an SAL of 10-6 for the complete STERRAD® 100S process. | DeCato K. and Chu K., 1996 RPT 01379 RI = 2 |
| Disinfection | Instruments and materials in hospitals and industry (PT2) | Hydrogen peroxide (59% nominal H2O2) (within the STERRAD® 50, total chamber volume: 73L) | G. stearothermophilus ATCC 7953 | Death kinetics 5 replicates per spore carrier material at each hydrogen peroxide volume tested. The carriers were placed in open glass Petri dishes placed on top of instruments in each of two trays. The validation load was then double wrapped with Heavy Duty Spunguard. Recovery methods: - plate count for volumes of 0.02, 0.045, 0.09 and 0.135 mL - direct culture for concentrations of 0.18, 0.27, 0.36, 0.72 and 1.44 mL. Controls: One positive (unprocessed, inoculated spore carrier, for each material type) and one negative control (media) were incubated each day testing was conducted. Bacteriostasis testing was performed for each material. | One half-cycle, with an injection volume of 0.02 to 1.44 mL of 59% hydrogen peroxide. Preconditioning = 15 min preconditioning plasma Half-cycle parameters set to: Injection = 6 min Diffusion = 2 min Plasma = 2 min Spore carriers (approx. 5 mm x 20 mm) inoculated with at least 1 x 106 G. stearothermophilus spores. Spore carriers were of the following material types: PTFE (PT) Polyurethane (PU) Polyethylene (PE) Silicone (S) Glass filters (GF) Stainless steel (SS) Monel (M) | Mean Log10 values for B. stearothermophilus. Spore survivors with increasing H2O2 injection volumes. H2O2 injection Vol. (µL); H2O2 conc. - chamber (mg/L); Material. SS; S; PU; M; GF; PT; PE. 20; 0.20; 4.04; 3.99; 3.25; 4.37; 3.80; 4.15; 3.64. 45; 0.45; 3.51; 2.64; 1.45; 3.45; 2.79; 3.81; 4.09. 90; 0.90; 1.53; 1.38; 1.81; 2.15; 1.51; 1.30; <1. 135; 1.34; 1.34; <1; <1; <1; 1.15; 1.41; 1.58. 180; 1.79; -0.04; -0.29; na; -0.04; -0.04; na; -0.29. 270; 2.68; -0.29; na; -0.29; na; na; na; na. 360; 3.58; na; na; na; na; na; na; na. 720; 7.16; na; na; na; na; na; na; na. 1440; 14.32; na; na; na; na; na; na; na. na – Log10 not calculated, no survivors Initial concentrations (log10 CFU per carrier): SS = 6.01, S = 6.19, PU = 6.06, M = 6.11, GF = 6.14, PT = 6.08, PE = 6.10 Positive controls : growth Negative controls : no growth No bacteriostatic effect from the process for any of the materials tested. Endospore death kinetics data obtained for each material tested demonstrates a positive “dose response” to increasing volumes of hydrogen peroxide injected at half cycle conditions. There were no spore survivors on any material with an injection volume of 360 µL or greater. The nominal Hydrogen peroxide volume for the half cycle is 1440 µL. The results obtained in this test, at a minimum, an SAL of 10-6 for the complete STERRAD® 50 process. | Roberts C., 1996 RPT 01315 RI = 2 |
| Disinfection | Instruments and materials in hospitals and industry (PT2) | Hydrogen peroxide (59% nominal H2O2) (within the STERRAD® 50 System, total chamber volume: 73L) | G. stearothermophilus ATCC 7953 | Half cycle validation and total kill end point validation – Stainless steel lumens (3 mm x 400 mm). One stainless steel spore coupon was placed at midpoint inside each of 10 stainless steel lumens. Lumens containing spore coupons were placed in the validation tray with medical instruments. The validation tray was double wrapped with Heavy Duty Spunguard wrap. The wrapped validation tray was placed on the shelf inside the STERRAD 50. One positive control (unprocessed spore carrier) and one negative control (media) were incubated each day testing was conducted. Bacteriostasis testing of the stainless steel spore carrier material was performed. | The load was processed at half cycle conditions : Preconditioning = 15 min preconditioning plasma Injection = 6 min Diffusion = 2 min Plasma = 2 min But injected volume of hydrogen peroxide was varied from 20 to 1440 µL. One set of 10 lumens for each condition was tested. The nominal half cycle (1440 µL of 59% nominal H2O2) tests were run in triplicate (3 separate cycles, each cycle containing 10 lumens). Upon completion of the sterilization process, spore carriers were aseptically transferred to separate tubes of TBS. The TBS tubes were incubated at 55°C for 21 days. | Volume of 59% nominal H2O2 injected (µL); H2O2 conc. - chamber (mg/L); Chamber pressure at end of injection (torr); Number of tubes positives N=10 (21 days). 20; 0.20; 1.37; 9/10 growth. 45; 0.45; 1.47; 10/10 growth. 90; 0.90; 1.81; 4/10 growth. 135; 1.34; 2.61; 1/10 growth. 180; 1.79; 3.22; 0/10 growth. 270; 2.68; 4.18; 0/10 growth. 360; 3.58; 5.24; 0/10 growth. 500; 4.97; 7.02; 0/10 growth. 720; 7.16; 9.11; 0/10 growth. 930; 9.25; 10.1; 0/10 growth. 1080; 10.74; 10.5; 0/10 growth. 1440; 14.32; 13.1; 0/10 growth. 1440; 14.32; 13.1; 0/10 growth. 1440; 14.32; 13.0; 0/10 growth. Repeat of small injection volume runs (20-180 µL) Volume of 59% nominal H2O2 injected (µL); H2O2 conc. - chamber (mg/L); Chamber pressure at end of injection (torr); Number of tubes positives N=10 (21 days). 20; 0.20; 1.32; 10/10 growth. 45; 0.45; 1.83; 7/10 growth. 90; 0.90; 2.28; 1/10 growth. 135; 1.34; 2.59; 0/10 growth. 180; 3.58; 3.39; 0/10 growth. Positive controls: growth Negative controls: no growth No bacteriostatic effect from the process for any of the materials tested. Results from Half-Cycle and total Kill End Point Validation tests demonstrate an SAL of 10-6 for the complete STERRAD® 50 process. | Roberts C., 1996 RPT 1320 RI = 2 |
| Disinfection | Instruments and materials in hospitals and industry (PT2) | Hydrogen peroxide (59% nominal H2O2) (within the STERRAD® 100S System) | G. stearothermophilus ATCC 7953 | Efficacy studies with substrate materials representative of the materials commonly used in reusable medical devices. 2 sterile glass petri dishes containing 5 each inoculated test carriers of designed material (5mm x 20 mm). The two trays in the validation load were double wrapped with Spunguard wrap placed on the shelf inside the STERRAD 100S. One positive control (unprocessed spore carrier) and one negative control (media) were incubated each day testing was conducted. | Inoculated spore carriers (with at least 1 x 106 G. stearothermophilus spores) were placed in open petri dishes on top of instruments in the validation load and processed at less than half cycle conditions: Preconditioning = 10 min preconditioning plasma Half-cycle parameters set to: Injection = 6 min Diffusion = 2 min Plasma = 2 min With reduced hydrogen peroxide (59% nominal H2O2) volumes injected : 720µL for all materials tested At higher volumes for materials showing growth at 720µL. Following exposure, each of the carriers was aseptically transferred to TSB and incubated at 55°C for 14 days. | 720µL injection volume of 59%H2O2 => 3.02 mg/L in the chamber Material; Cycle 1; Cycle 2; Cycle 3. Aluminum; 0/10; 0/10; 0/10. Brass; 0/10; 0/10; 0/10. Delrin; 0/10; 0/10; 0/10. Ethyl Vinyl Acetate (EVA); 0/10; 0/10; 0/10. Glass; 0/10; 0/10; 0/10. Glass fiber; 0/10; 0/10; 0/10. Kraton; 0/10; 0/10; 0/10. Monel; 0/10; 0/10; 0/10. Neopren; 10/10; n/t; n/t. Nylon; 4/10; 1/10; 3/10. PMMA; 0/10; 1/10; 1/10. Polycarbonate; 0/10; 0/10; 0/10. Polyethylene; 0/10; 0/10; 0/10. Polypropylene; 0/10; 0/10; 0/10. Polystyrene; 0/10; 0/10; 0/10. Polyurethane; 0/10; 1/10; 0/10. PVC; 0/10; 0/10; 1/10. Rubber, latex - Bandage; 9/10; 10/10; n/t. Rubber, latex - condom; 0/10; 1/10; 0/10. Rubber, latex - Penrose; 6/10; 10/10; n/t. Silicone - GE; 0/10; 0/10; 0/10. Silicone - MC; 0/10; 0/10; 0/10. Stainless Steel; 0/10; 0/10; 0/10. Teflon (PTFE); 0/10; 0/10; 2/10. n/t = not tested. If positives observed, then testing continued at next injection volume. 1080µL injection volume of 59%H2O2 => 4.53 mg/L in the chamber Material; Cycle 1; Cycle 2; Cycle 3. Neopren; 2/10; 1/10; n/t. Nylon; 0/10; 0/10; 0/10. PMMA; 0/10; 0/10; 0/10. Polyurethane; 0/10; 0/10; 0/10. PVC; 0/10; 0/10; 0/10. Rubber, latex - Bandage; 1/10; 2/10; n/t. Rubber, latex - condom; 0/10; 1/10; 0/10. Rubber, latex - Penrose; 0/10; 0/10; 0/10. Teflon (PTFE); 0/10; 0/10; 0/10. n/t = not tested. If positives observed, then testing continued at next injection volume. 1440µL injection volume of 59%H2O2: => 6.04 mg/L in the chamber Material; Cycle 1; Cycle 2; Cycle 3. Neopren; 0/10; 0/10; 0/10. Rubber, latex - Bandage; 0/10; 4/10; 1/10. Rubber, latex - condom; 0/10; 1/10; 0/10. Positive controls: growth Negative controls: no growth No bacteriostatic effect from the process for any of the materials tested. Spore log reduction was greater than or equal to 6 for all materials tested. | DeCato K. and Chu K., 1996 RPT 01387 RI = 2 |
| Disinfection | Instruments and materials in hospitals and industry (PT2) | Hydrogen peroxide (59% nominal H2O2) (within the STERRAD® 100S System) | G. stearothermophilus ATCC 7953 | Half cycle validation and total kill end point validation – Stainless steel lumens (3 mm x 400 mm). One stainless steel spore coupon was placed at midpoint inside each of 10 stainless steel lumens. Lumens containing spore coupons were placed in the validation tray with medical instruments. The validation tray was double wrapped with Heavy Duty Spunguard wrap. The wrapped validation tray was placed on the shelf inside the STERRAD 100S. One positive control (unprocessed spore carrier) and one negative control (media) were incubated each day testing was conducted. Bacteriostasis testing of the stainless steel spore carrier material was performed. | The load was processed at half cycle conditions : Preconditioning = 15 min preconditioning plasma Injection = 6 min Diffusion = 2 min Plasma = 2 min But injected volume of hydrogen peroxide was varied from 20 to 1440 µL. One set of 10 lumens for each condition was tested. The nominal half cycle (1440 µL of 59% nominal H2O2) tests were run in triplicate (3 separate cycles, each cycle containing 10 lumens). The cycles were run in each of 3 machines (triplicate) to provide a more comprehensive data set. Upon completion of the sterilization process, spore carriers were aseptically transferred to separate tubes of TBS. The TBS tubes were incubated at 55°C for 21 days. | Machine #022 Volume of 59% nominal H2O2 injected (µL); H2O2 concentration in the chamber mg/L; Chamber pressure at end of injection (torr); Number of tubes positives N=10 (21 days). 20; 0.08; 1.06; 10/10 growth. 45; 0.19; 1.16; 10/10 growth. 90; 0.38; 1.36; 10/10 growth. 135; 0.57; 1.49; 10/10 growth. 180; 0.76; 1.70; 6/10 growth. 270; 1.13; 2.60; 2/10 growth. 360; 1.51; 2.57; 0/10 growth. 500; 2.10; 3.25; 0/10 growth. 720; 3.02; 4.35; 0/10 growth. 930; 3.90; 5.29; 0/10 growth. 1080; 4.53; 6.27; 0/10 growth. 1440; 6.04; 7.60; 0/10 growth. 1440; 6.04; 7.64; 0/10 growth. 1440; 6.04; 7.65; 0/10 growth. Machine #138 Volume of 59% nominal H2O2 injected (µL); H2O2 concentration in the chamber mg/L; Chamber pressure at end of injection (torr); Number of tubes positives N=10 (21 days). 20; 0.08; 1.05; 10/10 growth. 45; 0.19; 1.05; 10/10 growth. 90; 0.38; 1.30; 10/10 growth. 135; 0.57; 1.38; 9/10 growth. 180; 0.76; 1.75; 7/10 growth. 270; 1.13; 2.09; 0/10 growth. 360; 1.51; 2.57; 0/10 growth. 500; 2.10; 3.25; 0/10 growth. 720; 3.02; 4.27; 0/10 growth. 930; 3.90; 5.33; 0/10 growth. 1080; 4.53; 6.06; 0/10 growth. 1440; 6.04; 7.34; 0/10 growth. 1440; 6.04; 7.47; 0/10 growth. 1440; 6.04; 7.35; 0/10 growth. Machine #139 Volume of 59% nominal H2O2 injected (µL); H2O2 concentration in the chamber mg/L; Chamber pressure at end of injection (torr); Number of tubes positives N=10 (21 days). 20; 0.08; 1.09; 10/10 growth. 45; 0.19; 1.02; 10/10 growth. 90; 0.38; 1.26; 10/10 growth. 135; 0.57; 1.43; 9/10 growth. 180; 0.76; 1.69; 7/10 growth. 270; 1.13; 2.08; 2/10 growth. 360; 1.51; 2.35; 1/10 growth. 500; 2.10; 3.21; 0/10 growth. 720; 3.02; 4.18; 0/10 growth. 930; 3.90; 5.11; 0/10 growth. 1080; 4.53; 5.96; 0/10 growth. 1440; 6.04; 7.32; 0/10 growth. 1440; 6.04; 7.49; 0/10 growth. 1440; 6.04; 7.41; 0/10 growth. Positive controls: growth Negative controls: no growth No bacteriostatic effect from the process for any of the materials tested. Each of the 3 machines produced comparable results. Results from Half-Cycle and total Kill End Point Validation tests demonstrate no spore survivors after the first half of the cycle at nominal hydrogen peroxide injection volumes (1440 µL) i.e. 6.04 mg/l in the chamber, thus an SAL of 10-6 for the complete STERRAD® 100S process. | DeCato K., 1996 RPT-01388 RI = 2 |
| Disinfection | Instruments and materials in hospitals and industry (PT2) | Hydrogen peroxide (59% nominal H2O2) (within the STERRAD® 100S System) | Bacillus subtilis ATCC 19659 Clostridium sporogenes ATCC 3584 | AOAC method 966.04 sporicidal activity of disinfectants using endospore contaminated silk suture loops and porcelain penicylinders were performed to evaluate the sterilization efficacy of the STERRAD 100S system. 3 tests were each performed with 120 B. subtilis carriers and 120 C. sporogenes carriers (60silk suture loops and 60 penicylinders) packaged in tyvek pouches. A total of 48 pouches were tested per cycle with each pouch containing 5 suture loops or 5 penicylinders. The pouches were placed on the racks in the sterilizer and processed through a full cycle. Bacteriostatic controls with non-contaminated carriers processed through the STERRAD 100S system were performed. | Full cycle conditions: Preconditioning = 10 min plasma Followed by 2 phase processes each with: 6 min injection 2 min diffusion 2 min plasma. Injection volume: 1440 µL of 59% nominal H2O2 => 6.04 mg/L in the chamber. Following exposure, the carriers were transferred to Fluid Thioglycollate Medium containing one drop of catalase reagent. At the completion of the first transfer, the carriers were re-transferred to fresh tubes of FTM and incubated at 37°C for 21 days. | Test results are expressed as numbers of growth after 21 days of incubation. B. subtilis; C. sporogenes. Sutures; Penicylinders; Sutures; Penicylinders. AOAC test. Test 1; 0/60; 0/60; 0/60; 0/60. Test 2; 0/60; 0/60; 0/60; 0/60. Test 3; 0/60; 0/60; 0/60; 0/60. Bacteriostatic control. Test 1; 2/2; 2/2; 2/2; 2/2. Test 2; 2/2; 2/2; 2/2; 2/2. Test 3; 2/2; 2/2; 2/2; 2/2. The STERRAD® 100S full cycle demonstrates sporicidal activity at 6.04 mg/L according to AOAC test method 966.04. | Chan-Myers H., 1996. RPT - 01426 RI = 2 |
| Disinfection | Instruments and materials in hospitals and industry (PT2) | Hydrogen peroxide (59% nominal H2O2) (within the STERRAD® 200 System) | B. subtilis ATCC 19659 C. sporogenes ATCC 3584 | AOAC method 966.04 sporicidal activity of disinfectants (same method as described above) | Full cycle conditions: Preconditioning = 10 min plasma Followed by 2 phase processes each with: 6 min injection 2 min diffusion 2 min plasma. Injection volume: 9.3 mg/L of 59% nominal H2O2 per half cycle. Carriers were then transferred to FTM tubes, each containing one drop of catalase reagent. At the completion of the first transfer, the carriers were re-transferred to fresh tubes of FTM. All tubes were then incubated at 37°C for 21 days. After 21 days of incubation, the FTM tubes with no growth results were heat-shocked for 20 min at 80°C and re-incubated for 72h at 37°C. | Test results are expressed as numbers of growth after 21 days of incubation at 37°C, 20 min of heat shock at 80°C and 72h of incubation at 37°C. B. subtilis; C. sporogenes. Sutures; Penicylinders; Sutures; Penicylinders. AOAC test. Test 1; 0/60; 0/60; 0/60; 0/60. Test 2; 0/60; 0/60; 0/60; 0/60. Test 3; 0/60; 0/60; 0/60; 0/60. Bacteriostatic control; 2/2; 2/2; 2/2; 2/2. The STERRAD® 200 full cycle demonstrates sporicidal activity at 9.3 mg/L according to AOAC test method 966.04. | Choi S., 1999. RPT-01821 RI = 2 |
| Disinfection | Instruments and materials in hospitals and industry (PT2) | Hydrogen peroxide (59% nominal H2O2) (within the STERRAD® 200 System) | G. stearothermophilus ATCC 7953 | Demonstration of effects on spore survival during various phases. One stainless steel spore coupon was placed at midpoint inside stainless steel lumens (3 mm x 400 mm). Lumens were then placed in the validation tray. The validation tray was double wrapped with Heavy Duty Spunguard wrap. The wrapped validation tray was placed on the shelf inside the STERRAD 200. The load was processed at different test conditions. A. Demonstration of pre-exposure plasma effects (enumeration of BI) B. Exposure 1 and Exposure 2 phases without pre-exposure plasma (direct culture of BI) C. Exposure 1 and Exposure 2 phases with pre-exposure plasma (direct culture of BI) | Full short cycle: Pre-exposure plasma: 20 min Followed by 2 phase processes Injection: 6.5 min Diffusion: 2 min Plasma: 2 min Application rates: - Adjusted amount of deionized water that mimics the same pressure rise as 9.3 mg/L of H2O2. - H2O2 : per half cycle, 9.3 mg/L. Biological controls : - positive control : unprocessed biological indicators (BI) - Negative control: Media Biological indicator: Spore carriers (stainless steel blade coupons) inoculated with at least 1 x 106 G. stearothermophilus spores. | A. Effect of pre-exposure plasma Test condition; Log10 Spore recovery; Spore Log Reduction. Unprocessed BI; 6.16; -. Pre-exposure Plasma only; 5.99; 0.17. Pre-exposure Plasma + exposure 1 (H2O); 6.07; 0.09. Pre-exposure Plasma + exposure 1 and 2 (H2O); 6.03; 0.13. B. Exposure 1 and Exposure 2 phases without Pre-exposure Plasma: Test conditions; Sterility testing results (# nonsterile / # tested). No pre-exposure plasma. Half cycle only. H2O2 injection.; 0/12, 0/12, 0/12. No pre-exposure plasma.1st half cycle, H2O injection; 2nd half cycle, H2O2 injection.; 0/12, 0/12, 0/12. Positive control; 2/2. Negative control; 0/2. C. Full cycles with pre-exposure Plasma: Test conditions; Biological results (# nonsterile / # tested). 1st half cycle, H2O2 injection. 2nd half cycle, H2O injection.; 0/12. 1st half cycle, H2O injection; 2nd half cycle, H2O2 injection.; 0/12. 1st half cycle, H2O2 injection. 2nd half cycle, H2O2 injection.; 0/12. 1st half cycle, H2O injection; 2nd half cycle, H2O injection.; 12/12. Positive control; 2/2. Negative control; 0/2. No bacteriostatic effect from the process for the material tested. The pre-exposure plasma phase of the STERRAD® 200 cycle results in very minimal antimicrobial activity. The complete STERRAD® 200 cycle delivers a SAL of 10-6 with and without pre-exposure plasma. Furthermore, this study demonstrates that a full cycle processed without H2O2 doesn’t lead to disinfection. | Chu N., 1999 RPT-01814 RI = 2 |
| Disinfection | Instruments and materials in hospitals and industry (PT2) | Hydrogen peroxide (59% nominal H2O2) (within the STERRAD® 200 System) | G. stearothermophilus ATCC 7953 | H2O2 dose response and validation of sterilization of lumens. Stainless steel coupons BIs were placed at the internal midpoint of 3 mm x 400 mm stainless steel lumens. The lumens with BIs were then placed in each of the 4 sterilization trays containing the STERRAD® 200 validation load. The trays were wrapped and placed inside the sterilizer chamber in specified locations and processed using the STERRAD® 200 at half cycle conditions with a pre-determined injection concentration of 59% nominal hydrogen peroxide. Recovery methods: - plate count for 0.93 and 1.88 mg/L - direct culture for all other concentrations. One positive control (unprocessed spore carrier) and one negative control (media) were incubated each day testing was conducted. Bacteriostasis testing of the stainless steel carrier material was performed. | Half cycle conditions: Pre-exposure plasma : 20 min Injection : 6 min Diffusion : 2 min Plasma : 2 min Application rates: incremental amounts ranging from 0.93 to 9.3 mg/L of 59% nominal H2O2 The nominal half cycle (9.3 mg/L of 59% nominal H2O2) tests were run in triplicate All samples were incubated at 55-60°C: 48h for plate count method, and 21 days for direct culture. | mg H2O2/L of chamber volume; Plate count method Spore log reduction (6 samples); Sterility test # pos. / # tested. 0.93; 2.04 to 3.10; -. 1.86; 4.99 to 6.31; -. 2.79; -; 8/12. 3.72; -; 3/12. 4.65; -; 0/12. 6.23; -; 0/12. 7.91; -; 0/12. 9.30; -; 0/12. 9.30; -; 0/12. 9.30; -; 0/12. Positive control; 2/2. Negative control; 0/2. Overall average titer of the positive control (4 samples): 2.02 x 106 CFU/BI No bacteriostatic effect from the process for the material tested. These results demonstrate that total kill was achieved, in a half cycle, within 3 mm x 400 mm stainless steel lumens when 4.65 mg/L of Hydrogen peroxide was injected into the chamber. An SAL of 10-6 was demonstrated in a full short STERRAD® 200 cycle. | Chan-Myers H., 1999 RPT-01902 RI = 2 |
| Disinfection | Instruments and materials in hospitals and industry (PT2) | Hydrogen peroxide (59% nominal H2O2) (within the STERRAD® System) | Range of microorganisms tested at 5% serum and 10% serum concentrations 5 replicates each | ISO 14937, Sterilization of healthcare products – General requirements for characterization of a sterilizing agent and the development, validation and routine control of a sterilization process AAMI Standards and Recommended Practices, Volume 1; Sterilization (1992) | STERRAD® Systems Abbreviated cycle consisting of 40 minutes of diffusion with 3 mg/L of H2O2 and 10 minutes of plasma at a power of 300 watts. | Microorganism; # Positive / # Tested. Pseudomonas aeruginosa; 0/10. Pseudomonas cepacia; 0/10. Xanthomonas maltophilia; 0/10. Serratia marcescens; 0/10. Klebsiella (encapsulated); 0/10. Methicillin-resistant Staphylococcus aureus; 0/10. Slime-producing Staphylococcus epidermidis; 0/10. Listeria monocytogenes; 0/10. Enterococcus faecalis; 0/10. Acinetobacter calcoaceticus; 0/10. Salmonella sp.; 0/10. Shigella sp.; 0/10. Campylobacter sp.; 0/10. Aeromonas sp.; 0/10. Clostridium perfringens; 0/10. Clostridium tetani; 0/10. Clostridium difficile; 0/10. Bacillus subtilis spores; 0/10. Micrococcus sp.; 0/10. Mycobacterium tuberculosis; 0/10. Mycobacterium chelonei; 0/10. Bacteroides fragilis; 0/10. Fusobacterium sp.; 0/10. Anaerobic cocci; 0/10. Candida albicans; 0/10. | Smith D. F., 2000* Supportive data No raw data |
| Disinfection | Instruments and materials in hospitals and industry (PT2) | Hydrogen peroxide (59% nominal H2O2) (within the STERRAD® 200 System) | Geobacillus stearothermophilus | Double-wrapped tray containing three mated surfaces and a standard load of medical devices. The mated surface biological challenge consisted of a material sample inoculated with 106 G. stearothermophilus spores mated with an identical material sample, such that the inoculum was sandwiched between the two mated surfaces. Four trays used per test. | Nominal half-cycle conditions in STERRAD® 200 | Validation test with mated surfaces: Test 1 Test 2 Test 3 #pos. / #pos. / #pos. / Material #tested #tested #tested Stainless steel 0/12 0/12 0/12 blades Polyacetal 0/12 0/12 0/12 (Delrin) Polyetherimide 0/12 0/12 0/12 (Ultem*) Polyarylsulfone 0/12 0/12 0/12 (Radel**) *Ultem is a trademark of GE Plastics **Radel is a trademark of Amoco | Smith D. F., 2000* Supportive data No raw data |
| Disinfection | Instruments and materials in hospitals and industry (PT2) | Hydrogen peroxide (58% w/w H2O2) (within the STERRAD® 100 System) | Pseudomonas aeruginosa (NCTC 6749), Mycobacterium tuberculosis (H37Ra), Clostridium perfringens (NCTC 8237, spores), Clostridium tetani (F7269/91, spores) | Instruments packed in plastic trays on a polypropylene wrap. A hydrogen peroxide adaptor was applied to the fibre optic colonoscope channel. Biological and chemical indicator strips included in each tray, which was then double-wrapped with polypropylene wraps and bound with special colour indicator tapes before processing in the sterilizer. | Full STERRAD® 100 Sterilization process (75 min). 6 mg/L H2O2 in the chamber. Inoculum of >106 on the following test pieces; orthopaedic flat metal French chisel, metal orthopaedic Chiro 1200 drill, Olympus fibre optic video colonoscope. Biological indicators containing >106 spores of Bacillus subtilis var. niger were also used in each test. | Total viable counts (log10 cfu/mL) on wash fluids after Sterrad® 100 process under clean and dirty conditions. Pa Mt Cp Ct Inoculum 6.18 4.56 6.00 6.54 Chisel 1st wash NG NG NG NG† 2nd wash NG NG NG NG† 3rd wash NG NG NG NG† Drill 1st wash NG NG NG NG† 2nd wash NG NG NG NG† 3rd wash NG NG NG NG† Scope 1st wash NG NG NG* NG† 2nd wash NG NG NG* NG† 3rd wash NG NG NG* NG† * C. perfringens was isolated from the colonoscope only, under clean condition when the adaptor attachment was omitted. † C. tetani was initially isolated on three occasions under dirty conditions. NG – no growth | Kyi M.S., Holton J. and Ridgway G.L., 1995* Supportive data No raw data |

* A synthesis of these references is presented in annex 3.5

Studies performed on Geobacillus stearothermophilus demonstrated the efficacy of the product HYDROGEN PEROXIDE SOLUTION, with conditions of use in compliance with those claimed (dose injected, number of cycles and H2O2 concentration in the chamber). The efficacy criterion of > 5 log reduction claimed is demonstrated.

As steps E-I are considered as a block and cannot be dissociated, efficacy data without the plasma cycle during this part is unavailable.

Some studies have been carried out in order to demonstrate that H2O2 is the active substance in the process. Study RP01814, shows that the pre-exposure plasma phase (step B, above), results in very minimal antimicrobial activity, and the desired > 5 log reduction cannot be achieved with the pre-exposure plasma step alone or even a full cycle processed without H2O2 (but with deionized water). So it confirms the function of the plasma stages, as described by the applicant (see above), i.e the degradation of peroxide to oxygen and water and reduction of H2O2 residuals on the articles.

Furthermore, studies RP-1388 and RP-1379, utilizing variable dosage of H2O2 injections in a “half cycle” (consisting of step A through I), shows microbial activity at low H2O2 dosage (two plasma phases: A and I) but cannot demonstrate the desired 6 log reduction. As the H2O2 dosage approaches approximately 25% of the nominal dosage of 1440 µL (6.04 mg/L in the chamber), the desired 6 log reduction is observed.

Based on the results of these studies, it can be concluded that disinfection at a level of > 5 log occurs at 6 mg H2O2/L or greater. So the claimed minimum application volume of 5 mL Hydrogen Peroxide Solution per injection stage (corresponding to 13.41 mg H2O2 /L in the chamber), with a minimum of 2 injection per cycle, is validated.

Conclusion on the efficacy of the product

French competent authorities (FR CA) assessed that the product HYDROGEN PEROXIDE SOLUTION, is efficient against bacteria, bacterial spores, yeasts, fungi and virus when used within the STERRAD 200 GMP system, to disinfect instruments and materials during manufacturing in the industrial sector (PT02), except the uses covered by the medical device regulation.

#### Occurrence of resistance and resistance management

According to the assessment report of the active substance: “the lethal effects of oxidative molecular species generated from hydrogen peroxide can be avoided with any damage being repaired in microorganisms such as Escherichia coli and Salmonella typhimurium. When E. coli and S. typhimurium are exposed to low concentrations of H2O2, 3 μM and 60 μM respectively, cells produce enzymes and other proteins which are important for cellular defence and mitigate the toxic effects of the oxidative species. This adaptive response is triggered by nontoxic levels of the oxidative species to protect against and produce resistance to oxidative stress caused when challenged with higher concentrations, 10 mM (Dukan and Touati (1996), Christman et al. (1985)). The resistance to oxidative stress that E.coli develops when exposed to H202, as reported in literature papers, demonstrates an adaptive response only. Hydrogen peroxide has been intensively used as a disinfectant and preservative for more than 3 decades and has not lead to the development of significant resistance levels among field populations. Genetically inherited resistance is not expected when the products are used as recommended”.

A study of Collinson and Dawes also reports resistance to H2O2 of the yeast Saccharomyces cerevisiae.

According to the applicant, no incidence of resistance to hydrogen peroxide has been recorded.

To ensure a satisfactory level of efficacy and avoid the development of resistance, the recommendations proposed in the SPC have to be implemented

#### Known limitations

According to the applicant, there are factors generally associated with product configuration, product packaging, and load configuration that can adversely influence sterilization.

Hydrogen peroxide plasma-based sterilization does require diffusion of the sterilant into the product and packaging to effect sterilization. Paper and other cellulosic-based materials are contraindicated for use in the STERRAD® Sterilizer because of high affinity of these kinds of materials for hydrogen peroxide. Powders are not recommended for sterilization in the STERRAD® Sterilizer because they limit diffusion of hydrogen peroxide. Liquids are also contraindicated due to the deep vacuum associated with this methodology.

There are materials that may act to degrade hydrogen peroxide when present in large amounts in the product load. These are materials such as copper, silver, platinum, palladium, and nickel alloys. It should be noted that small amounts of these metals, such as the amounts found in integrated circuits boards, are not normally detrimental to the functionality of the sterilizer.

Other materials that catalytically degrade hydrogen peroxide would also potentially negatively impact the hydrogen peroxide concentration. Large amounts of absorbtive materials may also adversely affect the sterilization process. Certain types of compounds, such as organic and polymeric disulfides will react with hydrogen peroxide to give potentially corrosive by-products. An example is the oxidation of molybdenum disulfide, a lubricant used in fiberoptic endoscopes, to create sulfuric acid. The sulfuric acid may etch polymeric materials with which it comes in contact. Additional information is found in the STERRAD System Operator’s Manual.

#### Evaluation of the label claims

French competent authorities (FR CA) assessed that the product HYDROGEN PEROXIDE SOLUTION, is efficient against bacteria, bacterial spores, yeasts, fungi and virus when used for disinfection within the STERRAD® 200 GMP System, at an injection volume of 5 mL (per injection stage) i.e. 13.47 mg H2O2 /L in the chamber and a sterilization cycle with minimum 2 injection stages, of instruments and materials during manufacturing process in industries (PT02), except uses covered by the medical device regulation.

Furthermore, sterilisation is not considered in the Biocide Regulation but only disinfection. The assessment of the product HYDROGEN PEROXIDE SOLUTION concludes on the effectiveness of this product as a disinfectant and not as a sterilant. Indeed, the purpose of sterilisation is of the responsibility of the user. The user shall define the routine monitoring and control procedures for their specific process.

#### Relevant information if the product is intended to be authorised for use with other biocidal product(s)

The product HYDROGEN PEROXIDE SOLUTION is not intended to be used with other biocidal products.

### Risk assessment for human health

The product HYDROGEN PEROXIDE SOLUTION is a biocidal product containing 59.8% w/w hydrogen peroxide. It is packaged in bottles of 500mL (HPDE).

The product is applied by vaporising hydrogen peroxide in the STERRAD® 200 GMP Sterilisation System, a low temperature sterilizer using the STERRAD process.

The product is intended for use by professional only for the sterilization of instruments and materials in hospitals and industry

#### Assessment of effects on Human Health

No acute toxicity study (oral, dermal and inhalation), nor skin and eye irritation study neither skin sensitisation study has been performed on the product.

Classification of the product has been carried out according to the calculation rules laid down in the CLP regulation.

Skin corrosion and irritation

| Data waiving |  |
| --- | --- |
| Information requirement | Skin corrosion and irritation |
| Justification | No study has been performed on the product HYDROGEN PEROXIDE SOLUTION. Regarding the content of active substance and co-formulant in the product, and according to the classification rules laid down in the CLP regulation, a classification Skin corrosion 1B, H314:“Causes severe skin burns and eye damage” is required. |

Eye irritation

| Data waiving |  |
| --- | --- |
| Information requirement | Eye irritation |
| Justification | No study has been performed on the product HYDROGEN PEROXIDE SOLUTION. Regarding the content of active substance and co-formulant in the product, and according to the classification rules laid down in the CLP regulation, a classification H314:“Causes severe skin burns and eye damage” is required. |

Respiratory tract irritation

| Data waiving |  |
| --- | --- |
| Information requirement | Respiratory tract irritation |
| Justification | No study has been performed on the product HYDROGEN PEROXIDE SOLUTION. Regarding the content of a.s and co-formulant in the product, a classification STOT-SE Category 3, H335:“ May cause respiratory irritation” should be required for respiratory tract irritation. Nevertheless, according to the classification rules laid down in the CLP regulation, classification as corrosive is considered to implicity cover the potential to cause respiratory tract irritation and so the addition of classification STOT-SE Category 3 H335 is considered to be superfluous. Considering the corrosive and acute toxicity (inhalation) classification of the product and according to the CLP Regulation, the label statement EUH071-“Corrosive to the respiratory tract” shall be applied. |

Skin sensitization

| Data waiving |  |
| --- | --- |
| Information requirement | Skin sensitization |
| Justification | No study has been performed on the product HYDROGEN PEROXIDE SOLUTION. Regarding the content of a.s and co-formulants, and according to the classification rules laid down in the CLP regulation, no classification is required for skin sensitization. |

Respiratory sensitization (ADS)

| Data waiving |  |
| --- | --- |
| Information requirement | Respiratory sensitization |
| Justification | No study has been performed on the product HYDROGEN PEROXIDE SOLUTION. Regarding the content of a.s and co-formulants, and according to the classification rules laid down in the CLP regulation, no classification is required for respiratory sensitization. |

Acute toxicity

The acute toxicity of hydrogen peroxide depends on its concentration.

In the CAR, the acute oral toxicity studies were performed with formulations containing hydrogen peroxide at concentrations from 35 % to 70 %, range containing the active substance concentration present in the product HYDROGEN PEROXIDE SOLUTION.

The acute inhalation toxicity studies were performed with formulations containing 49.3 % hydrogen peroxide, a concentration inferior to the active substance concentration in the product.

The acute dermal toxicity studies were performed with formulations containing 70 % hydrogen peroxide, concentration which represents a worst-case compared to the active substance concentration present in the product.

Acute toxicity by oral route

| Data waiving |  |
| --- | --- |
| Information requirement | Acute toxicity by oral route |
| Justification | No study has been performed on the product HYDROGEN PEROXIDE SOLUTION. Regarding the content of active substance and co-formulant in the product, and according to the classification rules laid down in the CLP regulation, an ETA mix of 712 mg/kg is obtained. Therefore, a classification Acute tox. 4, H302:“ Harmful if swallowed.” is required for acute oral toxicity. |

Acute toxicity by inhalation

| Data waiving |  |
| --- | --- |
| Information requirement | Acute toxicity by inhalation |
| Justification | No study has been performed on the product HYDROGEN PEROXIDE SOLUTION. According to the studies carried out in the CAR of active substance, the inhalation LC50 value for the test substance containing 49.3% hydrogen peroxide was > 0.17 mg/L/4 h (highest attainable vapour concentration). When corrected to 100% hydrogen peroxide, the LC50 value was > 0.084 mg/L. Regarding the content of active substance and co-formulant in the product, and according to the classification rules laid down in the CLP regulation, an ETA mix of 0.14 mg/L is obtained. Therefore, a classification Acute tox. 2, H330:“ Fatal if inhaled.” is required for acute inhalation toxicity. |

Acute toxicity by dermal route

| Data waiving |  |
| --- | --- |
| Information requirement | Acute toxicity by dermal route |
| Justification | No study has been performed on the product HYDROGEN PEROXIDE SOLUTION . Regarding the content of a.s and co-formulants, and according to the classification rules laid down in the CLP regulation, no classification is required for dermal acute toxicity. |

Information on dermal absorption

| Data waiving |  |
| --- | --- |
| Information requirement | Dermal absorption |
| Justification | No study has been performed on the product HYDROGEN PEROXIDE SOLUTION. Based on the corrosive property of hydrogen peroxide, 100% dermal absorption value should be used as default value. However, in the absence of clear systemic effects, no dermal absorption value is needed to conclude on human health risk assessment. |

Available toxicological data relating to non active substance(s) (i.e. substance(s) of concern)

According to the definition of a substance of concern laid down in the Guidance on the BPR Volume III Human Health – Part B and C Risk Assessment, HYDROGEN PEROXIDE SOLUTION does not contain any substance of concern.

Available toxicological data relating to a mixture

Not applicable.

Other

Not applicable.

#### Exposure assessment

The HYDROGEN PEROXIDE SOLUTION is used within the Sterrad® 200 GMP Sterilization System by professional to sterilize instruments and materials in hospitals and industry.

The biocidal product is marketed as bottles of hydrogen peroxide solution. As the bottle is designed for use solely within Sterrad® 200 GMP Sterilization System, dispensing of the hydrogen peroxide during each sterilization cycle will be controlled and consistent. The sterilization process takes place within the enclosed environment of the STERRAD system that is designed specifically to minimize any potential for release of hydrogen peroxide directly into the environment.

The process, which can be described in five phases, uses a gas-diffusion phase to penetrate all parts of the load before the plasma is created.

The five phases are:

- Vacuum phase: The chamber is evacuated, reducing internal pressure in preparation for the subsequent reaction;

- Injection phase: A measured amount of liquid peroxide is injected into the chamber, evaporating the aqueous hydrogen peroxide solution and dispersing it into the chamber, where it kills bacteria on any surface it can reach;

- Diffusion phase: The hydrogen peroxide vapour permeates the chamber, exposing all load surfaces to the sterilant and rapidly sterilizes devices and materials without leaving any toxic residues. At the completion of this phase, the chamber pressure is reduced and the plasma discharge is initiated;

- Plasma phase: An electromagnetic field is created in which the hydrogen peroxide vapour breaks apart, producing a low temperature plasma cloud that contains ultraviolet light and free radicals. Following the reaction, the activated components lose their high energy and recombine to form oxygen and water;

- Vent phase: The chamber is vented to equalize the pressure enabling the chamber door to be opened. There is no need for aeration or cool-down.

Phases 1, 2 and 3 are then run for a second time for added efficacy, reprocessing to assure optimal sterilization.

A plasma phase is run at the end of the cycle which converts residual hydrogen peroxide (loaded or on the sterilized product) to water vapour and molecular oxygen.

For the product, the active ingredient, hydrogen peroxide, is contained within the sealed bottle. The bottle has two caps: an outer cap (red cap) and an inner cap with a thin membrane in the middle (white cap)(figure 1). To load a new bottle, the operator must first pull of the red cap from the new bottle and then, turning it over, insert it in into the bottle carrier. The white cap should not be damaged (figure 2). Closing the lid to the bottle carrier, a sharp blade at the bottom of the carrier punctures the membrane inside the bottle (white cap) and the hydrogen peroxide fills an accumulator inside the sterilizer.

Figure 1: Bottle of hydrogen peroxide with two caps.

Figure 2: Bottle carrier and insertion of the bottle into the carrier.

When the bottle is empty, it is automatically ejected into a collection box. When the collection box is full, the system alerts the user to discard the collection box.

The system is designed to specifically minimize any potential contacts with the active substance. There is no handling of open bottles. No mixing of chemicals or connecting of tanks is required. The treatment area is not accessible during the sterilization process.

In theory any residual hydrogen peroxide is expected to react quickly with organic matter on the surfaces, leading to its decomposition thereby limiting indirect dermal exposure to hydrogen peroxide.

In practice there is potential for exposure to residual hydrogen peroxide vapour by inhalation route during and after the sterilization process (by the evacuation system and during opening of the door of The Sterrad® system chamber) and by dermal route upon handling sterilized materials.

Hydrogen peroxide solution is used by professionals in a closed system and the treatment area is not accessible during the sterilization process. So, no indirect exposure of the general public, as a result of application is expected.

Identification of main paths of human exposure towards active substance and substances of concern from its use in biocidal product

| Summary table: relevant paths of human exposure |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Exposure path | Primary (direct) exposure |  |  | Secondary (indirect) exposure |  |  |  |
|  | Industrial use | Professional use | Non-professional use | Industrial use | Professional use | General public | Via food |
| Inhalation | n.a | Yes | n.a | n.a | Yes | n.a | n.a |
| Dermal | n.a | Yes | n.a | n.a | Yes | Yes | n.a |
| Oral | n.a | No | n.a | n.a | No | n.a | n.a |

List of scenarios

| Summary table: scenarios |  |  |  |
| --- | --- | --- | --- |
| Scenario number | Scenario (e.g. mixing/ loading) | Primary or secondary exposure Description of scenario | Exposed group (e.g. professionals, non-professionals, bystanders) |
| 1. | Inhalation of residual hydrogen peroxide vapor in air | Residual vapour of hydrogen peroxide emitted by Sterrad® 200 GMP Sterilization system during the process and/or when opening the door could be inhaled by the operator | Professional |
| 2. | Handling of sterilised material | After the sterilization process, there is potential for exposure to residual hydrogen peroxide when the sterilized material is handled. | Professional |
| 3. | Contact with sterilised material | After the sterilization process, there is potential of exposure to residual hydrogen peroxide when the sterilized material is in contact withtthepatient. | General Public |

Professional exposure

Scenario [1]: inhalation of residual hydrogen peroxide vapor in air

The notifier provided a monitoring study of measurement of Hydrogen peroxide emissions from a STERRAD® 200 System. Both 200 GMP and STERRAD 200 utilizes the STERRAD process, where H2O2 is used as the main sterilant, supplemented by plasma phases to and remove residual H2O2 from the sterilizer emission.

The measurements were made following OSHA ID-006 Sampling and Analytical Procedure which measures concentrations of hydrogen peroxide emissions at specified locations (fully validated method). In this way, a spectrophotometric method developed for air monitoring by OSHA in which hydrogen peroxide vapor in air reacts with titanium oxosulfate to form a colored solution was used.

During this monitoring, air samples are continually collected and measured from a single sterilizer (STERRAD® 200 Sterilizer, Serial Number 10202100039) operation period to simulate occupational exposure over a typical 8 hour workday (at a sampling rate of 0.5 L/Min.).

The OSHA method recommends only 200 minutes, but to be conservative, samples are collected over approximate 5 short cycles of 75 minutes (375 minutes). The interval between cycles is approximately 25-35 minutes. This interval simulates usage by a customer, where this represents opening of the machine door (10 min), load removal, cleaning, load reassembly and transport.

However, air samples were collected for a total of 10-11 hours while the five cycles runs were being performed. The extra sampling time is accounted for because of a need to re-run a cycle caused by an error loading the consumable cassette into the sterilizer. During the 10-11 hour period, the titanium oxysulfate solution was replaced with fresh solution midway through the sample time.

Air samples were collected at three target locations:

- at the door in front of the STERRAD 200 (1,5 m above the floor and 45 cm (18 inches) directly in front of the machine) samples STM-1A, STM-1B (figure 3)

- at the fan exhaust on top of the machine (45 cm (18 inches) above exhaust opening)  samples STM-2A, STM-2B (figure 4)

- in the laboratory at a distance of 3 to 4.5 meters (10 to 15 feet) from the STERRAD 200 machine being tested.  samples STM-3A, STM-3B (figure 5)

Figure 3: Sampling location in front of STERRAD 200 unit, 1,5 m above the floor and 45 cm  directly in front of the machine

Figure 4: Sampling location, on top of STERRAD 200 machine, 45 cm above exhaust opening

Figure 5: Sampling location, 4.5 meters directly in front of the STERRAD 200 machine

Both samples at each of three locations were submitted for analysis.

Results of the monitoring are summarized in Table 1 below:

Table1: Results of area air sampling and analysis for hydrogen peroxide

| Sample No. | Sample Description | Sample duration (minutes) | Results (ppm) | Results (mg/m3) |
| --- | --- | --- | --- | --- |
| STM-1A | STERRAD 200 – In front of the unit | 214 | <0.21 | <0.29 |
| STM-1B | STERRAD 200 – In front of the unit | 433 | 0.10 | 0.14 |
| TWA |  | 647 | <0.14 |  |
| STM-2A | STERRAD 200 – At exhaust opening at the top of the unit | 217 | <0.21 | <0.30 |
| STM-2B | STERRAD 200 – At exhaust opening at the top of the unit | 440 | <0.11 | <0.15 |
| TWA |  | 657 | <0.14 |  |
| STM-3A | Inside the second floor of the R&D Laboratory (15 feet in front from the STERRAD machine being tested) | 214 | <0.22 | <0.31 |
| STM-3B | Inside the second floor of the R&D Laboratory (15 feet in front from the STERRAD machine being tested) | 447 | <0.11 | <0.15 |
| TWA |  | 661 | <0.15 |  |
|  | TWA: Time weighted average <: Less Than |  |  |  |

In Table 1, for each sample, the quantity of hydrogen peroxide measured is under of the limit of quantification (32 µg). Results expressed as mg/m3 vary due to variation in the sample time and volume collected.

Since the quantity of hydrogen peroxide measured is under of the limit of quantification for each sample, results given are an estimate of the probable hydrogen peroxide concentration present in the measured sample. These results are maximizing against the actual concentration.

The worst case is obtained for the sample STM-3A with value lower than 0.31 mg/m3.

According FR CA, this monitoring study presents several deficiencies:

- emission  of H2O2 is measured from STERRAD 200 System  and not  200 GMP

- as the measure has been done in ambient air and not on the operator (which would have been preferred), the sampling rate of 0.5L/min is low to determine the exposure of worker, notably to detect a peak of exposure. A sampling rate of 2L/min should have been used.

- the loading of product is performed by loading of cassette and not bottle, which does not correspond to the use of HYDROGEN PEROXIDE SOLUTION

- the removal of box with the empty bottles is not taken in consideration

As a consequence, this study is not accepted by eCA for the estimation of primary exposure.

Scenario [2]: Use of sterilised materials

In practice there is potential exposure to residual hydrogen peroxide after the sterilization process upon handling materials.

However, the plasma phase is expected to reduce residual hydrogen peroxide on the process load. This step degrades peroxide to oxygen and water and reduces residuals on the product sterilized. Furthermore, any residual hydrogen peroxide is expected to react quickly with organic matter on the surfaces, leading to its decomposition thereby limiting indirect dermal exposure to hydrogen peroxide.

The applicant provided two studies to determine the level of hydrogen peroxide residual remaining on materials.

A first study (Scenario 2a) was submitted in order to measure the hydrogen peroxide residuals remained on surface of different materials (polymeric and metal formulations) after a Sterrad® process. (Technical Dossier LC-10200-701 - Appendix 13).

For that purpose, 101 sterilisation cycles were performed and a standard extraction of material in deionized water was employed, followed by a liquid chemical titration of resulting extract. For materials, at the minimum, 1 g of sample is tested (surface area ≥1 cm2) H2O2 residuals were reported in ppm on a weight basis and mg H2O2/cm2 of material.

Results indicate that for some materials no residual hydrogen peroxide is detectable after sterilization processes (e.g. polypropylene, silicone, teflon, alluminum...). But for some materials as nylons or polyurethane, the residual hydrogen peroxyde on the material reaches 10000 ppm that corresponds to 10 mg of H2O2/cm2 (in a worst-case calculation).

In this study, the number of performed cycle is very high and far from real uses conditions. The measured level of residual hydrogen peroxide could be induced by the high number performed cycle in absence of organic material, which could be a not realistic situation.

Another more realistic study was provided (Scenario 2b), in which the level of residual hydrogen peroxide remaining on medical devices (colonofiberscope, resectoscope set and penrose drain) after sterilization in the STERRAD 200 Sterilization cycle was determined (Technical Dossier LC-10200-701 - Appendix 03). This study will be used to characterize the risk.

Table 2: Composition of medical devices

| Device | Material composition |  |  |  |
| --- | --- | --- | --- | --- |
|  | Stainless Steel | Other materials | Glass | Plastics |
| Penrose drains |  |  |  | x |
| Colonofiberscope | x | x | x | x |
| Resectoscope set |  | x | x | x |

The test materials were chosen to represent the range of hydrogen peroxide interactivity of polymeric materials and metals that are included in the sterilizer claims for medical devices. The colonofiberscope was chosen because of the large outer surface of its insertion tube composed primarily of polyurethane that is known for high absorptivity of H2O2.

The cycle configuration employed was, maximum hydrogen peroxide injection, full cycle, minimum temperature settings and low plasma power, to represent the worst case exposure in the sterilizer processing.

The full STERRAD 200 cycle was performed 3 times for each tested devices and samples was taken in triplicate. Aliquots from irrigation channel of colonofberscope was also taken. For materials, at the minimum, 1 g of sample is tested (surface area ≥1 cm2) H2O2 residuals were reported in ppm on a weight basis (table 3).

Table 3: Results of measured residual hydrogen peroxide on the device

| Device | Mean [H2O2] (ppm) | Standard Deviation |
| --- | --- | --- |
| Colonofiberscope irrigation channel | 0.36 | 0.27 |
| Colonofiberscope | 148 | 4.31 |
| Resectoscope set | 13 | 1.48 |
| Penrose drains | 429 | 29.88 |

Results indicate that:

- The residual levels determined in the colonofiberscope irrigation liquid were negligible (0.36+/-0.27 ppm)

- In the other devices, the level of hydrogen peroxide residual is low and the highest value is achieved on the penrose drain (429+/- 29.9 ppm) or 0.429 mg of H2O2/cm2 (in a worst-case calculation).

Calculations for Scenario [2]

| Summary table: exposure from professional uses |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Exposure scenario | Tier/PPE | Estimated inhalation exposure | Estimated dermal exposure | Estimated oral exposure | Estimated total exposure |
| Scenario [2] | Tier 1/ No PPE |  | 0.429 mg of H2O2/cm2 |  | 0.429 mg of H2O2/cm2 |

Non-professional exposure

Not relevant

Exposure of the general public

Scenario [3]: Contact with sterilized material

The exposure of the general public in contact with sterilised material is covered by the exposure of professionals in contact with sterilised material.

Please refer to Scenario [2].

Monitoring data

For details, please refer to exposure descriptions above.

Exposure associated with production, formulation and disposal of the biocidal product

Not relevant

Aggregated exposure

Not relevant

Summary of exposure assessment

| Scenarios and values to be used in risk assessment |  |  |  |
| --- | --- | --- | --- |
| Scenario number | Exposed group (e.g. professionals, non-professionals, bystanders) | Tier/PPE | Estimated total exposure |
| 1. | professionals | Tier 1/ No PPE | 0.31 mg/m3 |
| 2. | professionals | Tier 1/ No PPE | 0.429 mg of H2O2/cm2 |
| 3. | general public | Tier 1/ No PPE | 0.429 mg of H2O2/cm2 |

Dietary exposure

By definition PT2 biocidal product is for application on surfaces that are not used for direct contact with food or feeding stuffs. Therefore residue in food or feed are not expected.

Information of non-biocidal use of the active substance

| Summary table of other (non-biocidal) uses |  |  |  |
| --- | --- | --- | --- |
|  | Sector of use1 | Intended use | Reference value(s) 2 |
| 1. | Plant protection product | hydrogen peroxide (basic substance – approved on 29/03/2017) | No MRLs required (Reg 396/2005) |
| 2. | Veterinary use | Hydrogen peroxide: all food producing species | No MRL required (Reg 37/2010) |

1 e.g. plant protection products, veterinary use, food or feed additives

2 e.g. MRLs. Use footnotes for references.

Estimating Livestock Exposure to Active Substances used in Biocidal Products

Not relevant

Estimating transfer of biocidal active substances into foods as a result of professional and/or industrial application(s)

Not relevant

Estimating transfer of biocidal active substances into foods as a result of non-professional use

Not relevant

#### Risk characterisation for human health

In the CAR of the active substance hydrogen peroxide, no systemic effect have been observed and no systemic reference toxicological value has been derived. Only a local effect risk assessment has been performed.

Reference values to be used in Risk Characterisation

| Reference | Study | NOAEL (LOAEL) | Safety Factor | Correction for oral absorption | Value |
| --- | --- | --- | --- | --- | --- |
| AEC inhalation short-term | NOAEC in 90-day inhalation study (rat) | - | 8 | - | 1.25 mg/m3 |
| AEC inhalation medium-term | NOAEC in 90-day inhalation study (rat) | - | 8 | - | 1.25 mg/m3 |
| AEC inhalation long-term | NOAEC in 90-day inhalation study (rat) | - | 8 | - | 1.25 mg/m3 |
| ARfD | Not established |  |  |  |  |
| ADI | ADI not established, the substance is not systemically available. The agreed acceptable max concentration is 0.1 mg/L in human drinking water |  |  |  |  |

Risk for professional users

Scenario [1]: inhalation of residual hydrogen peroxide vapor in air

The applicant provided a monitoring study of measurement of hydrogen peroxide emissions from a STERRAD® 200 System. The highest concentration measured is for the sample STM-3A with value lower than 0.31 mg/m3. This value has to be compared compared with the AECinhalation of hydrogen peroxide  equal to 1.25 mg/m3.

However, considering the deficiencies of the exposure study, FR CA considers that exposure of the operator below  AECinhalation cannot be fully guaranteed. Therefore the RMM has to be implemented :

“The user has to wear a hydrogen peroxide gas detector. If the threshold of 1.25 mg / m3 is exceeded, the room has to be evacuated.”

Dermal local risk assessment

A qualitative local risk assessment according to the guidance on the BPR: Volume III HH part B is performed since the product is corrosive (Cat1B).

| Hazard |  |  | Exposure |  |  |  |  |  |  | Risk |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Hazard Category | Effects in terms of C&L | Additional relevant hazard information | PT | Who is exposed? | Tasks, uses, processes | Potential exposure route | Frequency and duration of potential exposure | Potential degree of exposure | Relevant RMM & PPE | Conclusion on risk |
| High hazard | Skin Corr. 1B | - | 2 | Professional | Loading of bottle | Dermal and inhalation | 1/day Few minutes per day | Negligible as to load a new bottle, the operator must first pull of the red cap from the new bottle and then, turning it over, insert it in into the bottle carrier. The white cap should not be damaged | Measures to ensure well controlled exposure, such as: Technics - Minimisation of manual phases; - Regular cleaning of equipment and work area; - Avoidance of contact with contaminated tools and objects; Organisation - Minimise number of staff exposed; PPEs: Substance/task appropriate gloves;. | According to the guidance for concluding qualitatively on the acceptability for professional exposure, practically no exposure and no splashes should occur to lead acceptable risk. Exposure would be comparable to brief contact as touching of contaminated surface. Considering that exposure will be negligible because the system is automated and closed, the risk is considered acceptable with the proposed PPEs (gloves). |

Scenario [2]: Use of treated device

The levels of hydrogen peroxide residual remaining on devices (2b) after disinfection in the STERRAD 200 system in more realistic conditions (Technical Dossier LC-10200-701 - Appendix 03) were low and the highest value is achieved on the penrose drain (with 0.429 mg of H2O2/cm2). Considering a transfer of 100% from device to the skin (worst-case condition), a person will be exposed at a maximum of 0.429 mg of H2O2/cm2.

In the CAR of the hydrogen peroxide active substance, there is a skin irritation study performed with a 10% (w/w) solution in rabbits (OCDE 404)(Cf. Doc IIIA of CAR). In this study, 0.5 mL of solution is applied to an area of 6 cm2 during 4h. These conditions correspond to an exposure of the skin to 0.8 mg of  H2O2/cm2. No skin irritation is observed.

The exposure described in the skin irritation study (0.8 mg of H2O2/cm2) is superior to the maximum exposure which could be induced by contact with disinfected devices (0.429 mg of H2O2/cm2). So, the hydrogen peroxide remaining on the device does not present any hazard for the human health.

Combined scenarios

Not relevant because of local effects.

Conclusion for professional user

The risk for professional users is acceptable.

Conclusion

The risk for professional user is acceptable.

Risk for non-professional users

Not relevant

Risk for the general public

Scenario [3]: Contact with sterilised device

The risk for the patient in contact with sterilised devices is the same than professionals in contact with treated devices.

The hydrogen peroxide remaining on the device doesn’t present any hazard for the human health.

Conclusion on general public

The risk for general public is acceptable.

Risk characterisation from combined exposure to several active substances or substances of concern within a biocidal product

Not relevant.

Maximum residue limits or equivalent

Not relevant

Risk for consumers via residues in food

By definition PT2 biocidal product is for application on surfaces that are not used for direct contact with food or feeding stuffs. Therefore residue in food or feed are not expected.

General Conclusion on human health

The risk for the professional using the product HYDROGEN PEROXIDE PRODUCT within the STERRAD® 200 GMP Sterilization System, is acceptable when gloves are worn during the handling of product and when the following RMM is followed: “The user has to wear a hydrogen peroxide gas detector. If the threshold of 1.25 mg / m3 is exceeded, the room has to be evacuated. The risk for professional users and general public in contact with treated device is acceptable.

### Risk assessment for animal health

Not relevant

### Risk assessment for the environment

Please notice that the risk assessment for the environment (section 2.2.9) is reported as provided by the applicant. The FR CA position is presented in green evaluation boxes.

#### Effects assessment on the environment

No new data on the product is submitted. The product composition is similar to the representative product reviewed for inclusion of the active substance on the Union list. The intended use is broadly also similar to the one of the representative uses for the inclusion of the active substance in the Union list of approved substances under Regulation No. 528/2012.

Hydrogen peroxide decomposes rapidly in different environmental compartments. The

following processes are involved in the decomposition/degradation of hydrogen peroxide in the environment:

o	Biotic degradation catalysed by microbial catalase and peroxidase enzymes

o	Abiotic degradation by:

o	transition metal (Fe, Mn, Cu) and heavy metal catalysed decomposition

o	oxidation or reduction reactions with organic compounds or formation of addition compounds with organic or inorganic substances

Hydrogen peroxide decomposes into water and oxygen (2H2O2 → 2H2O + O2). The rate of this reaction depends on the contact with catalytic materials and other factors such as heat and sunlight. Hydrogen peroxide shows a very rapid biodegradation in sewage sludge with a DT50 of 2 minutes (at 20°C). Rapid degradation of hydrogen peroxide has also been observed in surface water and soil compartments. This degradation has been proposed to be mainly microbially derived.

The biotic and abiotic decomposition reactions proceed in parallel with the formation reactions and the equilibrium of these reactions depends on the environmental conditions. The measured/estimated representative half-lives for different environmental compartments are:

Air – 24 hours

Surface water – 5 days

Soil – 12 hours

Sewage sludge – 2 minutes

The low measured value of Henry's law constant of (H = 7.5 × 10-4 Pa × m3/mol) indicates very low volatilisation of hydrogen peroxide from water. As hydrogen peroxide is miscible with water in all proportions and taking into account that the calculated log KOC is 0.2036 mL/g, it is expected that hydrogen peroxide has a low potential for adsorption to soil and for partitioning to suspended matter or sediment. The estimated log KOW of –1.57 indicates negligible potential of bioconcentration of hydrogen peroxide in biota. The BCFs calculated according to TGD for fish and earthworm are 1.4 and 0.84, respectively. Therefore, no accumulation of hydrogen peroxide in the food chain is expected.

PBT and POP assessment

Hydrogen peroxide is rapidly degraded in a sewage sludge, surface water and soil. Therefore hydrogen peroxide does not fulfil the criteria for persistent compound. The estimated log KOW is -1.57 and the BCF for fish and earthworm are 1.4 and 0.84, respectively. therefore the bioaccumulation criterion is not fulfilled.

Hydrogen peroxide is an inorganic compound, therefore it does not fulfil the criteria for persistent organic pollutant (POP). The vapour pressue of hydrogen peroxide is 214 Pa and the estimated atmospheric half-life is 24 hours (less than 2 days); therefore the criteria for POP as defined in Annex D of the Stockholm Convention is not fulfilled.

The long-term NOEC value for Daphnia magna for hydrogen peroxide is 0.64 mg/L, and hydrogen peroxide is not classified for CMR properties and for specific target organ toxicity after repeated exposure. Therefore the toxic criterion is not fulfilled. As hydrogen peroxide does not fulfil all the 3 criteria, it is not a PBT substance.

Infobox 1 - FR CA position: We agree with data presented above. The active substance Hydrogen peroxide is not classified according to the harmonised classification. Nevertheless, Hydrogen peroxide should be classified H 412 according to the available data of the CAR. Indeed, Daphnia was the most sensitive aquatic organism with the lowest chronic ecotoxicity endpoint (21d): NOEC= 0.63 mg/L and the substance is considered as rapidly degradable. Therefore, the classification of the product HYDROGEN PEROXYDE SOLUTION is aquatic chronic category 3: H412. PNECs; CAR. STP microorganims; 4.66 mg/L. freshwater; 12.6E-3 mg/L. Freshwater sediments; 0.0101 mg.kgwwt-1 (equilibrium). soil; 1.84E-03 mg.kgwwt-1 (equilibrium).

#### Exposure assessment

Emission estimation

Environmental exposure assessment

The STERRAD system is designed so that direct exposure to hydrogen peroxide is unlikely. The most likely source of exposure of the environment is via accidental spillage.

Emission to air is unlikely since vapourised hydrogen peroxide is decomposed within the STERRAD system after the decontamination process. There are no tubes or pipes connecting the STERRAD system to discharge systems , therefore direct discharge into the sewage system will not occur.

Indirect (accidental) contamination of the environment is the most likely exposure route for the environment. In the competent authority report, the predicted environmental concentration (PEC) in STP effluents for PT2 was 0.0031 µg/L. The worst-case predicted environmental concentration in sewage effluents was 31 µg/L for PT4. After degradation in the STP, residual hydrogen peroxide may reach surface water, and the PECs in surface water receiving sewage treatment effluents were in the range of 0.00031 to 3.1 µg/L. The PEC in surface water is below typical background levels in surface water.

Hydrogen peroxide in surface water does not partition to suspended matter or sediment. The worst-case calculated PEC in sediment for all intended uses assessed was 0.0013 mg/kg.

Emissions of hydrogen peroxide to air from the STERRAD system will not occur, therefore calculation of PEC in air is not necessary.

No direct emission to soil from the STERRAD system is expected, therefore calculation of PEC in soil is not necessary.

In the risk assessment we have used a risk envelope approach based on the data evaluated during the review for inclusion into the Union list of authorised substances. The risk characterisation ratio (RCR) for microorganisms in a sewage treatment plant and surface water for all intended uses evaluated is below 1, indicating hydrogen peroxide does not cause unacceptable risk to biological processes at the sewage treatment plant and to aquatic organisms and sediment-dwelling organisms. The risk assessment for the intended use of hydrogen peroxide is considered to be within this envelope.

Infobox 2 - FR CA position: The HYDROGEN PEROXIDE SOLUTION is contained within a sealed bottle, which is not opened at any time but is inserted in the closed STERRAD® 200 GMP Sterilizer. The product used to sterilize equipment is applied by vaporisation in the closed system STERRAD® 200 GMP Sterilisation System. The instructions of use are clearly defined by the applicant. This system is designed to minimize any potential for releases of hydrogen peroxide directly or indirectly into the environment. In fact, after a diffusion phase of hydrogen peroxide in the chamber to expose all surfaces and rapidly sterilize devices and materials, a plasma phase is run at the end of the cycle which converts hydrogen peroxide to water vapour and molecular oxygen. So, the plasma phase reduces residual hydrogen peroxide in the chamber and on the sterilized materials to negligible level. Furthermore, the STERRAD® 200 GMP Sterilizer is not connected with tubes or pipes to discharge systems; consequently direct release of the product HYDROGEN PEROXIDE SOLUTION into the sewage system will not occur. Given the expected total decomposition of H2O2 following the use of the STERRAD® 200 GMP Sterilisation System, and the lack of connection between the sterilizer and the sewage system, no emissions into the environmental compartments (STP, water, sediment, soil and groundwater) are expected with the normal functioning of the STERRAD system. In this context, the environmental assessment for the product HYDROGEN PEROXIDE SOLUTION is not relevant.

Primary and secondary poisoning

Potential for secondary poisoning

The estimated log KOW of hydrogen peroxide is –1.57 indicating a negligible potential for bioconcentration in biota. Therefore, accumulation of hydrogen peroxide in the food chain is not expected, and the risk of secondary poisoning in aquatic and terrestrial predators is considered negligible.

Hydrogen peroxide is not included in the Commission staff working document on implementation of the Community Strategy for Endocrine Disrupters - a range of substances suspected of interfering with the hormone systems of humans and wildlife (COM (1999) 706)). There is no evidence of any endocrine disruption potential in the human health or ecotoxicological studies presented in the dossier.

Infobox 3 - FR CA position: Primary and secondary poisoning assessment for the product HYDROGEN PEROXIDE SOLUTION is not relevant. No ED properties are foreseen for the active substance or the co-formulants of this product.

#### Risk characterisation

Infobox 4 - FR CA position: No emissions are expected into the environment; consequently no quantitative environmental risk assessment is considered related to the use of the product HYDROGEN PEROXIDE SOLUTION.

Infobox 5 - FR CA position: To conclude, no emissions are expected into the environment; consequently no quantitative environmental risk assessment is considered related to the use of the product HYDROGEN PEROXIDE SOLUTION. However, only the use of HYDROGEN PEROXIDE SOLUTION in the closed STERRAD® 200 GMP Sterilizer leads to acceptable risks to the environment. Therefore, the following risk mitigation measure can be proposed: “Use only in the STERRAD system without tubes or pipes connecting to sewage systems”

### Measures to protect man, animals and the environment

[Please refer to summary of the product assessment and to the relevant sections of the assessment report.]

### Assessment of a combination of biocidal products

Not relevant

### Comparative assessment

Not relevant

# Annexes

## List of studies for the biocidal product

| Section no, / Reference no. | Author(s) | Citation in Style of EU risk Assessment Report | Year | Title. Source (where different from company) Company, Report no. GLP (where relevant) / (un)published | Data Protection claimed (Yes/No) | Owner |
| --- | --- | --- | --- | --- | --- | --- |
| 2.2.5.5 | Smith DF | Not applicable | 2000 | STERRAD® 200 STERILIZATION SYSTEM | Y | ASP |
| 2.2.5.5 | Kyi MS., et al. | Kyi et al. (1995) | 1995 | Assessment of the efficacy of a low temperature hydrogen peroxide gas plasma sterilization system. Journal of Hospital Infection (1995) 31, 275-284 | N |  |
| 2.2.5.5 | Roberts C., Antonoplos P. |  | 1998 | Inactivation of human immunodeficiency virus type 1, hepatitis A virus, respiratory syncytial virus, vaccinia virus, herpes simplex virus type 1, and poliovirus type 2 by hydrogen peroxide gas plasma sterilization. AJIC Am J Infect Control 1998; 26: 94-101 | N |  |
| 2.2.5.5 |  |  |  | STERRAD 200 Sterilizer technical dossier for compliance with ISO 14937. Ref. 99046 |  |  |

## Output tables from exposure assessment tools

## New information on the active substance

Not relevant

## Residue behaviour

Not relevant

## Summaries of the efficacy studies

Not relevant

## Confidential annex

Please see the confidential annex

Please fill in here the identifying product name from R4BP.

Jacobs, P.T Mechanism and validation Issues in Gas Plasma Sterilization, in Frontiers in Sterilization Practice, The future of low-Temperature Technology, ed. Robins, D.S., Communicore, Newport Beach, CA. pp. 73-84.

Sterrad 100 hydrogen peroxide gas plasma sterilization system 510(k) application. US FDA K915682.

Krebs M.C., Bécasse P., Verjat D., Darbort J.C. Gas-plasma sterilization: relative efficacy of the hydrogen peroxide phase compared with that of the plasma phase. International Journal of Pharmaceutics. 1998. 160:75-81.

Block, Seymour. Peroxygen compounds. Pp. 167-181. In Disinfection, Sterilization and Preservation. Fourth edition, 1991. Seymour Block editor. Lea & Ferbiger.

Kokubo M., Inoue T. and Akers J. Resistance of common Environmental Spores of the genus Bacillus to Vapor Hydrogen Peroxide. 1998. PDA Journal of Pharmaceutical Science & Technology, 52:228-231.

EU Pharmacopoeia 27.3 Sec. 4-1-2

FDA Guidance: Biological Indicator (BI) Premarket Notification [510(k)] Submissions; October 4, 2007

Collinson L.P., Dawes I.W., Inductibility of the response of yeast cells to peroxide stress. Journal of General Microbiology (1992), 138, pp. 329-335.

Guidance on the BPR: Volume III human health - part B Risk assessment

If an IUCLID file is not available, please indicate here the summaries of the efficacy studies.